## Slide 1: Instrument characterization discussion
Characterization of Picarro system
Characterization of closed-path Licor systems (hybrid 7500 and the 7200)
Other instruments…

## Slide 2: Picarro
Picarro G1301
based on Wavelength-Scanned Cavity RingDown Spectroscopy (WS-CRDS)
Precision (@5sec)
CO2 <0.2 ppmv
CH4 <1 ppbv
H2O 100 ppmv
More stable
The “perfect” instrument?
LI-7500A
based on non-dispersive infrared (NDIR) measurements
Precision (@10Hz)
CO2 0.11 ppmv
H2O 0.0047 pptv
spectral interferences and motion problems.
LI-7200
Licor

## Slide 3
To limit at 10% the contribution of sensor noise to uncertainty in flux measurement by eddy covariance :
Rowe, M. D., Fairall, C. W., and Perlinger, J. A.: Chemical sensor resolution requirements for near-surface measurements of turbulent fluxes, Atmos. Chem. Phys. Discuss., 10, 24409-24433, doi:10.5194/acpd-10-24409-2010, 2010.
              with
R sensor resolution requirement
αc dimensionless solubility of C
water-air difference in concentration
kc air-water transfer velocity of scalar C [m s-1]
Sensor resolution requirements?

## Slide 4
Example: at a height of 18 m, wind speed 8 m s-1, u* = 0.28 m s-1, fNyquist = 5 Hz, and neutral stability (z/L = 0).
APcov' = 161 for covariance
If ΔCDMS = 2.0 ppbv R = 0.058 ppbv
If ΔCCO2 = 20 ppmv R = 0.054 ppmv
If ΔCCO2 = 100 ppmv R = 0.27 ppmv
 | CPc | R/ΔC
CO2 | 1.7E-04 | 2.7E-03
DMS | 1.8E-03 | 2.9E-02
Acetone | 2.0E-02 | 3.2E-01
SO2 | 3.5E-03 | 5.6E-02
CO | 5.6E-06 | 9.0E-05
O3 | 6.3E-05 | 1.0E-03

## Slide 5: Other things to consider…
Open path & Closed-path:
Response time
Sensitivities
Closed-path:
Flow in turbulent regime (Re>2300)
Lag time (use of a ‘puff system’)
Residence time in measurement cell
Attenuation effects due to sample line, cell, dryer, …
Water vapor effects: WPL vs Nafion membrane dryer?
Inlet filters or regular line cleanings?

## Slide 6: Example of Picarro instrument issues…
Evaluation of CO2 cospectra from CalNex. Can the instrument in it's current configuration see the CO2 flux when pCO2 is large at moderate wind speeds?
2) Water vapor correction.
poor performance of the Picarro H2O measurement  need to evaluate the effect of a Nafion dryer on the CO2 signal.
Hopefully we can remove the water signal and avoid the need for a correction.

## Slide 7
3) Investigation of the high frequency response characteristics. What causes the strange pattern in the high frequency variance spectrum? Is motion an issue? If so, is the motion affecting pressure control, optical elements, or something else?

## Slide 8: And with Chris’ magic (example time series with large sea-air temperature and humidity differences so that measurements of temperature and humidity fluctuations well above instrument noise levels)
w-c cospectra versus frequency: raw Licor w-c cospectrum (blue), the Picarro w-c cospectrum (dashed), corrected spectrum as per formula below is green; the w-( cm’-µq’) cospectrum is red.
CO2 variance spectra.
Corrected Licor spectra agree with the Picarro spectrum.

## Slide 9: Objectives
Tests to optimize the instruments performance:
In the laboratory: response time, sensitivities, frequency response, motion simulation…
Outside test like at the Table Mountain research facility north of Boulder, CO for instance. Purpose: significant H2O/CO2 flux for frequency response study
Integration of the specific instruments into flux system
Deployment and testing during cruises (real motion simulation).
Analysis of fluxes and more.

## Slide 10
SW DMS:
UH will build a smaller version of their DMS APIMS dedicated to sea water measurements.
Already have a small mass spectrometer for this purpose
Byron currently working on the design modifications necessary for APIMS conversion.
It would be desirable to get a collaborative cruise with someone (Steve Archer?) using the traditional GC measurement for DMS as validation for the new measurement.
CO:
Need to evaluate the performance of the LGR CO instrument for CO flux.
Need to put together a seawater measurement for CO.
Equilibrator:
Needed for seawater measurements of CO2 and CO.
Existing technology and not too expensive to build (what do we need to develop system? Requirements?)
We can address this early next year and test on the first cruise.
Need to buy or borrow a closed path Licor for the pCO2 measurement.
Other: Acetone, SO2, etc… ?
Other topics