## General
Surface Flux Team Daily Checks
If nothing to report use NR
Explanation
General
Date: | 20191027 | [YYYYMMDD]
Operator1: | Matt Shupe, Dave Costa | [name]
Operator2: | Ola Persson | [name]
Status of Operators: | NR | Any concerns or limitations regarding operators? Health, injury, etc.
General Weather Conditions: | Mostly clear and cold. Light breeze. | Brief weather synopsis, especially weather that would impact operations
General Sea Ice Conditions: | No change | Brief local sea ice synopsis, especially conditions that would impact operations
Special Operations: | Sunday. Limited activity on the ice. | Any daily activities or special operations that would possibly impact daily schedule or operations
Command Center
Utilities: | NR | [comments or NR] Status on power connection, power panel, heaters, lights, network connection
Physical Integrity: | NR | [comments or NR] Any damage, ingress of water/snow, snow/water management, other structural concerns
DAQ Computer: | NR
Workstation Computer: | NR
Communications Mode(s): | Radio, Iridum, and Cable | [Radio, Iridium, Cable, and/or None] Which comms systems are being used?
Communications Quality: | NR | Any information about operation of the radio and/or Iridium modems. What is the quality of connection? Any actions that were taken to modify the quality
Remote Mounts
Iridium Antenna - ARM Container | NR | Physical system assessment. Describe any actions taken such as cleaning or adjustments
Radio modem/antenna(s) - Crow's Nest | Lots of rime on the starboard radio antenna on the Crow's nest, but the port side antenna is clean. We are considering a way to mitigate rime build up. | Physical system assessment. Describe any actions taken such as cleaning or adjustments
Time-Lapse Camera - Crow's Nest | NA | Operational? Comment on visual assessment, and any actions taken. Data downloaded?
ACTION ITEMS
| Set up Finnish radiometer suite, calibrate Max's VISSS | List things that need to be accomplished in the coming days, developing issues, etc.
Supplies
| Schuko to US adapters, Tie wraps | Ongoing list of supplies needed in the field; carry this forward until supplies have been obtained

## Met Tower
Met Towers at Met City (11m tower and 30m mast)
Explanation
Local Conditions
Ice and Snow Conditions | No major change in conditions. | Ice thickness, cracks, snow drifting, ablation or any other relevant conditions for operations.
Ice and Snow Management | None | Anything required for physical maintenance of local area
Utility Status
Power line connectivity | up and stable | [up or down, comment on stability]
Network connectivity | up and stable | [up or down, comment on stability]
Utility cables | Good, but need to integrate the WXT-30m cable in with the rest of the bundle, as well as Blomquist network cables (to be run). | [good or in need of maintenance]. Utility cables on poles or maintained routinely on surface, connections secure, strain relief intact, etc.
Control box | NR | Any comments on the operations of the computer, data logger, power distribution, or other components. Temperature range and stability inside.
SPC Control box | NR | Any comments on Frey's power and control box.
Structural Assessment - 11m
Tower Base | Stable and level. | Stability of the base, including any potential ablation
Guy Anchors | Solid | Any sign of anchor instability?
Guy Tightness | Even | Even tightness? If not adjust.
Structural Integrity | Not problems identified | Any issues identified with structural components? Bolts, clamps, cables checked for tightness?
Plumbness | The tower is leaning to the North (towards the bear guard outpost on the ridge) and away from the location with the missing guy line. We are considering mitigation strategies. | Assess any shifts in plumbness of tower, possibly using manual measurements and/or sonic inclinometers
Boom Level / Stability | Reasonable level, but some tilt consistent with the overall tower tilt. | Assess stability and level of booms
Cabling | Solid | Secure with no rubbing.
Structural Assessment - 30m
Tower Base | NR | Stability of the base, including any potential ablation
Guy Anchors | NR | Any sign of anchor instability?
Guy Tightness | Tight on upwind side and slack on downwind side | Even tightness? If not adjust.
Structural Integrity | Not problems identified | Any issues identified with structural components? Bolts, clamps, cables checked for tightness?
Plumbness | Tower has slight un-evenness in the vertical and a lean with the wind. | Assess any shifts in plumbness of tower, possibly using manual measurements and/or sonic inclinometers
Cabling | Good, but need to deal with better mounting the instrument cabling at the base and making short power poles for the cabling heading back towards Met Hut. | Secure with no rubbing.
Instrument/Measurement Assessment
Daily Cleaning | 00:45-1:00 UTC cleaning of tower to remove heavy frosting (~1") on instruments and tower structure. | Include the timing and any details related to cleaning
PTU_2m | Cleaned frost off of the housing | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment, and any actions taken.
HMT_6m | Cleaned frost off of the housing | "
HMT_11m | Rime build up, but did not clean | "
WXT_30m | NR | "
Sonic_2m | Clear of rime | "
Sonic_6m | Clear of rime | "
Sonic_11m | Clear of rime | "
Sonic_30m | Ian and Ola working on orientation relative to the sonics on the 11m tower. | "
Licor | Cleaned frost off of the system with feather duster. Even after this there was still a red blinking light indicating an error likely due to frosting of the window. | "
SR50 | System was already clear of frost; measured snow depths of 1.5, 1.5, 3, 3 cm (avg 2.3 cm) underneath; SR50 distance was 187.9 cm at this time | "
IRT | Cleaned frost off of the system | "
Flux Plates | snow measurements made near each. For Plate A S of crack, measurements were 9, 9, 9, 9, 9, 6, 10 cm (9 cm avg); for Plate B, see SR50 assessment above | "
CLASP | Cleaned frost off of the system | "
SPCs (multiple) | Appear to be operational, but still working to get the data archival functional. | "
Data Files
Clasp02m |  | hourly files with size 3,095kB
CR1000X |  | daily file with size ~22500kB,
Licor02m |  | hourly files, raw with size ~3700-4000kB, stats with size ~19kB
Metek02m |  | hourly files, raw with size ~5280kB, stats with size 25kb
Metek06m |  | hourly files, raw with size ~5280kB, stats with size 25kb
Metek10m |  | hourly files, raw with size ~5280kB, stats with size 25kb
Metek30m |  | hourly files, raw with size ~5280kB, stats with size 25kb
Ozone |  | ????
Picarro |  | hourly files with size 3586kB

## ASFS40 (L1)
Atmospheric Surface Flux Stations in the Distributed Network
Note: Use this form for daily, remote assessments of ASFS. Use dedicated ASFS Site Visit form for site visits
Explanation
Communication System Status
Comms System | radio | [Radio, Iridium, or None] Indicate which forms of communication have been active
Quality | good | Provide any information on the quality of connection, how this impacts data flow
Power System Status
Operational Status | good | [Good, Questionable, Unknown]
Assessment of Exhaust | NR | Any notes about the buildup of ice that may be visible from the camera
Fuel Available | NR | Current remaining fuel available
Comments | EFOY power profile similar to recent, good 24V, 12V power is solid | Any other assessment of the power system, including quality of battery voltage/charge, operation of EFOY, etc.
Conditions
Structural | NR | Any useful information on the structural integrity of the ASFS
Position/Orientation/Level | heading ~350deg, both inclinometer readings about 0.5 to 0.6 deg, | Any useful information on position; i.e., rapid changes in GPS position, GPS orientation, or sonic-inclinometer level
Surface | NR | Any useful information on local surface conditions from camera
Sky-Weather | NR | Any useful information on weather or sky conditions
Instrument/Measurement Assessment
LWD | Consistent with mostly clear early, then moving towards ice clouds and then eventually low thick clouds / fog and/ or possibly some radiometer riming. fan good | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment from imagery if possible.
LWU | consistent with LWD, fan good | "
SWD | Very small values, but these also decrease when the clouds get thicker? fan and heat good | "
SWU | consistent with SWD, fan and heat good | "
IRT | Values are consistently about 0.2C colder than T_2m with a slight delay of a couple minutes in timing | "
SR50 | see temperature signal in the data, but also possibly a decrease in distance (increase in surface) due to some snowfall and/or drifting in the higher winds. | "
PTU | T_2m warms when clouds come and LWD increases, RH increases at the same time | "
Sonic | reasonable | "
Licor | NA | "
Flux Plates | Tracking with an offset ranging from 6-8 W/m2 over the day | "
Camera_Rad | NR | Any pictures available?
Camera_EFOY | NR | Any pictures available?
Data Logger | mainscan ~3800, Ptemp=-5 to 1C | Max scan time. Reasonable temperature range? Fans working?
Data Files
ASFS_40_slow |  | daily files of ~1120kB
ASFS_40_met |  | daily files of ~1740kB
ASFS_40_fast |  | daily files of ~5280kB
ASFS_40_Iridium_sumfile | NA | accumulating file with time stamp [include time stamp]

## ASFS30 (L2)
Atmospheric Surface Flux Stations in the Distributed Network
Note: Use this form for daily, remote assessments of ASFS. Use dedicated ASFS Site Visit form for site visits
Explanation
Communication System Status
Comms System | radio | [Radio, Iridium, or None] Indicate which forms of communication have been active
Quality | good | Provide any information on the quality of connection, how this impacts data flow
Power System Status
Operational Status | good | [Good, Questionable, Unknown]
Assessment of Exhaust | NR | Any notes about the buildup of ice that may be visible from the camera
Fuel Available | NR | Current remaining fuel available
Comments | EFOY power profile is similar to 10/24 | Any other assessment of the power system, including quality of battery voltage/charge, operation of EFOY, etc.
Conditions
Structural | NR | Any useful information on the structural integrity of the ASFS
Position/Orientation/Level | inclinometer continues to show very slow change in X tilt | Any useful information on position; i.e., rapid changes in GPS position, GPS orientation, or sonic-inclinometer level
Surface | NR | Any useful information on local surface conditions from camera
Sky-Weather | NR | Any useful information on weather or sky conditions
Instrument/Measurement Assessment
LWD | Follows LWU very closely, indicative of full rime coverage of the system. | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment from imagery if possible.
LWU | Looks reasonable | "
SWD | ~0 | "
SWU | ~0 | "
IRT | generally follows the Vaisala T-2m very closely with only up to about 0.2 C difference. | "
SR50 | Distance getting larger? | "
PTU | All looks reasonable | "
Sonic | reasonable | "
Licor | NA | "
Flux Plates | reasonable | "
Camera_Rad | NR | Any pictures available?
Camera_EFOY | NR | Any pictures available?
Data Logger | Main scan <4000, Ptemp= -5 to -2C | Max scan time. Reasonable temperature range? Fans working?
Data Files
ASFS_30_slow | good | daily files of ~1120kB
ASFS_30_met | good | daily files of ~1740kB
ASFS_30_fast |  | daily files of ~5280kB
ASFS_30_Iridium_sumfile | NA | accumulating file with time stamp [include time stamp]

## ASFS50 (L3)
Atmospheric Surface Flux Stations in the Distributed Network
Note: Use this form for daily, remote assessments of ASFS. Use dedicated ASFS Site Visit form for site visits
Explanation
Communication System Status
Comms System | iridium | [Radio, Iridium, or None] Indicate which forms of communication have been active
Quality | good enough | Provide any information on the quality of connection, how this impacts data flow
Power System Status
Operational Status | good | [Good, Questionable, Unknown]
Assessment of Exhaust | NR | Any notes about the buildup of ice that may be visible from the camera
Fuel Available | NR | Current remaining fuel available
Comments | Information in files is unclear | Any other assessment of the power system, including quality of battery voltage/charge, operation of EFOY, etc.
Conditions
Structural | NR | Any useful information on the structural integrity of the ASFS
Position/Orientation/Level | heading steady, no major change in sonic inclination | Any useful information on position; i.e., rapid changes in GPS position, GPS orientation, or sonic-inclinometer level
Surface | NR | Any useful information on local surface conditions from camera
Sky-Weather | NR | Any useful information on weather or sky conditions
Instrument/Measurement Assessment
LWD | consistent with mostly clear day | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment from imagery if possible.
LWU | consistent with LWD | "
SWD | NA | "
SWU | NA | "
IRT | continued close tracking with Vaisala T | "
SR50 | clear temperature sensitivity in the SR50 reading | "
PTU | reasonable | "
Sonic | reasonable | "
Licor | NA | "
Flux Plates | Reasonable but two plates of opposite sign. | "
Camera_Rad | NR | Any pictures available?
Camera_EFOY | NR | Any pictures available?
Data Logger | main counts ~120, Ptemp = -4C | Max scan time. Reasonable temperature range? Fans working?
Data Files
ASFS_50_slow |  | daily files of ~1120kB
ASFS_50_met |  | daily files of ~1740kB
ASFS_50_fast |  | daily files of ~5280kB
ASFS_50_Iridium_sumfile |  | accumulating file with time stamp [include time stamp]

## ARM
ARM Facility - General Data Oversight
Explanation
Measurement Assessments |  | Assessment of data for scientific quality. Do the measurements look reasonable compared to conditions? Are the measurements capturing the key features?
Clouds
KAZR | Vagner fixed a connection in the waveguide and this resolved an issue I had noticed with some horizontal lines in the reflectivities over the past days.
MWACR
Ka-SACR
X-SACR
HSRL
MPL
Ceilometer
MWR2C
MWR3C
TSI
Precipitation
LDIS/PARS2 (ice)
LDIS/PARS2 (ship)
PWD (ice)
PWD (ship)
Siphon rain gauge (ship)
ORG (ice)
WBRG/Pluvio (ice)
Winds
BSRWP
DL
Radiation
AERI
ShipRad - PIR, PSP, SPN1
SkyRad - PSP, PIR, SPN1 (ice) | domes are clean at 0430 UTC
GndRad - PIR, PSP (ice) | sensors look clean from a distance at 0430 UTC
MFR (ice)
CSPHOT
FRSR
Aerosol
AOSMet
CO / Purge
O3
CPC
CPCU
SMPS
HTDMA
UHSAS
CCNC
ACSM
SP2
PSAP
Nephelometer
INP filters
Data Requests |  | Any facilitation of MOSAiC data users

## Guests
Guest and Partner Instruments
Explanation
Power Distribution Unit - Met Hut (PFS)
Operational status | Operational and healthy.
Power balance on two PDUs? | Good balance, although still not much usage
Back-up generator test (/month) | NR
Precip VISSS - Met City (Maahn)
Cleaned rime off of light housing at about 00:45 UTC
Doppler Lidar - Met City (Brooks)
Doppler Lidar - Polarstern (Brooks)
Doppler Lidar - Polarstern (Heinemann)
Sodar - Met City (Brooks)
Radiation Suite - Met City (Pirazzini)
Not yet installed