## General
Surface Flux Team Daily Checks
If nothing to report use NR
Explanation
General
Date: | 20191029 | [YYYYMMDD]
Operator1: | Matt Shupe, Dave Costa | [name]
Operator2: | Ola Persson | [name]
Status of Operators: | NR | Any concerns or limitations regarding operators? Health, injury, etc.
General Weather Conditions: | Cloud coverage, no additional rime buildup on instrumentation | Brief weather synopsis, especially weather that would impact operations
General Sea Ice Conditions: | No change | Brief local sea ice synopsis, especially conditions that would impact operations
Special Operations: | Met City visit 9-11:30am local | Any daily activities or special operations that would possibly impact daily schedule or operations
Command Center
Utilities: | NR | [comments or NR] Status on power connection, power panel, heaters, lights, network connection
Physical Integrity: | NR | [comments or NR] Any damage, ingress of water/snow, snow/water management, other structural concerns
DAQ Computer: | NR
Workstation Computer: | NR
Communications Mode(s): | Radio, Iridum, and Cable | [Radio, Iridium, Cable, and/or None] Which comms systems are being used?
Communications Quality: | NR | Any information about operation of the radio and/or Iridium modems. What is the quality of connection? Any actions that were taken to modify the quality
Remote Mounts
Iridium Antenna - ARM Container | NR | Physical system assessment. Describe any actions taken such as cleaning or adjustments
Radio modem/antenna(s) - Crow's Nest | Lots of rime on the starboard radio antenna on the Crow's nest, but the port side antenna is clean. We are considering a way to mitigate rime build up. | Physical system assessment. Describe any actions taken such as cleaning or adjustments
Time-Lapse Camera - Crow's Nest | NA | Operational? Comment on visual assessment, and any actions taken. Data downloaded?
ACTION ITEMS
| Set up Finnish radiometer suite, calibrate Max's VISSS | List things that need to be accomplished in the coming days, developing issues, etc.
Supplies
| Schuko to US adapters, Tie wraps | Ongoing list of supplies needed in the field; carry this forward until supplies have been obtained

## Met Tower
Met Towers at Met City (11m tower and 30m mast)
Explanation
Local Conditions
Ice and Snow Conditions | No major change in conditions, some local drifting around the met tower base | Ice thickness, cracks, snow drifting, ablation or any other relevant conditions for operations.
Ice and Snow Management | None | Anything required for physical maintenance of local area
Utility Status
Power line connectivity | up and stable | [up or down, comment on stability]
Network connectivity | up and stable | [up or down, comment on stability]
Utility cables | Good, but need to integrate the WXT-30m cable in with the rest of the bundle, as well as Blomquist network cables (to be run). | [good or in need of maintenance]. Utility cables on poles or maintained routinely on surface, connections secure, strain relief intact, etc.
Control box | Opened box shortly today to attempt install of Blomquist network cable, although it was too short. Saw brief decrease in box temperature but then moving back towards stabilization. | Any comments on the operations of the computer, data logger, power distribution, or other components. Temperature range and stability inside.
SPC Control box | NR | Any comments on Frey's power and control box.
Structural Assessment - 11m
Tower Base | Stable and level. | Stability of the base, including any potential ablation
Guy Anchors | Solid | Any sign of anchor instability?
Guy Tightness | Even | Even tightness? If not adjust.
Structural Integrity | Not problems identified | Any issues identified with structural components? Bolts, clamps, cables checked for tightness?
Plumbness | The tower is leaning to the North (towards the bear guard outpost on the ridge) and away from the location with the missing guy line. Today we attempted to hammer a wooden shim wedge under the tower base plate, but this did not bring about any changes to the overall tower plumbness. | Assess any shifts in plumbness of tower, possibly using manual measurements and/or sonic inclinometers
Boom Level / Stability | Reasonable level, but some tilt consistent with the overall tower tilt. | Assess stability and level of booms
Cabling | Solid | Secure with no rubbing.
Structural Assessment - 30m
Tower Base | NR | Stability of the base, including any potential ablation
Guy Anchors | NR | Any sign of anchor instability?
Guy Tightness | NR | Even tightness? If not adjust.
Structural Integrity | Not problems identified | Any issues identified with structural components? Bolts, clamps, cables checked for tightness?
Plumbness | NR | Assess any shifts in plumbness of tower, possibly using manual measurements and/or sonic inclinometers
Cabling | NR, short power poles were installed to route the cables | Secure with no rubbing.
Instrument/Measurement Assessment
Daily Cleaning | 01:15-01:30 UTC Climbed tower for some basic maintenance. This included removing frost from the upper tower structure to minimize buildup over time, also removing old frost from the 10m met sensors. Importantly used some ethanol to remove a very fine layer of frost on the licor window, which was not visible from the surface. Also removed the utility rope from the tower. The climb was successful and safe, but non-trivial. | Include the timing and any details related to cleaning
PTU_2m | No additional frost/rime present | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment, and any actions taken.
HMT_6m | No additional frost/rime present | "
HMT_11m | Removed frost | "
WXT_30m | NR | "
Sonic_2m | Clear of rime; 0046 UTC, adjusted sonic orientation slightly clockwise to align better with tower "north" | "
Sonic_6m | Clear of rime | "
Sonic_11m | Clear of rime | "
Sonic_30m | NR | "
Licor | Cleaned frost off of the system using ethanol. It was not possible to see from the surface, but could be seen at the height of the licor. After cleaning this off, the blinking red light became solid again. | "
SR50 | System was already clear of frost. Surface under SR50 changed slightly due to some slight drifting around the cable for the flux plate installed in that location | "
IRT | No additional frost/rime present | "
Flux Plates | Slight drifting and scouring around the cables | "
CLASP | No additional frost/rime present | "
SPCs (multiple) | Appear to be operational, but still working to get the data archival functional. Removed some frosting from the arms of the 10m sensor | "
Data Files
Clasp02m |  | hourly files with size 3,095kB
CR1000X |  | daily file with size ~22500kB,
Licor02m |  | hourly files, raw with size ~3700-4000kB, stats with size ~19kB
Metek02m |  | hourly files, raw with size ~5280kB, stats with size 25kb
Metek06m |  | hourly files, raw with size ~5280kB, stats with size 25kb
Metek10m |  | hourly files, raw with size ~5280kB, stats with size 25kb
Metek30m |  | hourly files, raw with size ~5280kB, stats with size 25kb
Ozone |  | ????
Picarro |  | hourly files with size 3586kB

## ASFS40 (L1)
Atmospheric Surface Flux Stations in the Distributed Network
Note: Use this form for daily, remote assessments of ASFS. Use dedicated ASFS Site Visit form for site visits
Explanation
Communication System Status
Comms System | Radio | [Radio, Iridium, or None] Indicate which forms of communication have been active
Quality | Signal appeared to improve substantially at 5:08 UTC, going from 200 to 24 attempts, which appears to be no failed communication attempts. | Provide any information on the quality of connection, how this impacts data flow
Power System Status
Operational Status | Good | [Good, Questionable, Unknown]
Assessment of Exhaust | NR | Any notes about the buildup of ice that may be visible from the camera
Fuel Available | ?? | Current remaining fuel available
Comments | voltages look good both on the 24 and 12V sides. | Any other assessment of the power system, including quality of battery voltage/charge, operation of EFOY, etc.
Conditions
Structural | NR | Any useful information on the structural integrity of the ASFS
Position/Orientation/Level | NR | Any useful information on position; i.e., rapid changes in GPS position, GPS orientation, or sonic-inclinometer level
Surface | NR | Any useful information on local surface conditions from camera
Sky-Weather | NR | Any useful information on weather or sky conditions
Instrument/Measurement Assessment
LWD | consistent with cloudy for about 1/2 the day and clear for the other half Good fan | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment from imagery if possible.
LWU | consistent with LWD. Good fan | "
SWD | near 0, good fan and heat | "
SWU | near 0, good fan and heat | "
IRT | typically within about 0.1C of the vaisala T-2m | "
SR50 | continue to see strong temperature signal | "
PTU | reasonable | "
Sonic | reasonable | "
Licor | NA | "
Flux Plates | Follow eachother well, with offset of about 6 W/m2 | "
Camera_Rad | NR | Any pictures available?
Camera_EFOY | NR | Any pictures available?
Data Logger | max scan = 3400, logger temperature ~ 6C | Max scan time. Reasonable temperature range? Fans working?
Data Files
ASFS_40_slow |  | daily files of ~1120kB
ASFS_40_met |  | daily files of ~1740kB
ASFS_40_fast |  | daily files of ~5280kB
ASFS_40_Iridium_sumfile | NA | accumulating file with time stamp [include time stamp]

## ASFS30 (L2)
Atmospheric Surface Flux Stations in the Distributed Network
Note: Use this form for daily, remote assessments of ASFS. Use dedicated ASFS Site Visit form for site visits
Explanation
Communication System Status
Comms System | radio | [Radio, Iridium, or None] Indicate which forms of communication have been active
Quality | Radio comms quality improved at about 5:08 UTC, this coordinated improvement suggests that the change in the radios was due to something on our end. | Provide any information on the quality of connection, how this impacts data flow
Power System Status
Operational Status | good | [Good, Questionable, Unknown]
Assessment of Exhaust | NR | Any notes about the buildup of ice that may be visible from the camera
Fuel Available | NR | Current remaining fuel available
Comments | EFOY power profile is similar to 10/24 | Any other assessment of the power system, including quality of battery voltage/charge, operation of EFOY, etc.
Conditions
Structural | NR | Any useful information on the structural integrity of the ASFS
Position/Orientation/Level | Tilt looks the same, no rapid changes | Any useful information on position; i.e., rapid changes in GPS position, GPS orientation, or sonic-inclinometer level
Surface | NR | Any useful information on local surface conditions from camera
Sky-Weather | NR | Any useful information on weather or sky conditions
Instrument/Measurement Assessment
LWD | Follows LWU very closely, indicative of full rime coverage of the system. | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment from imagery if possible.
LWU | Looks reasonable | "
SWD | ~0 | "
SWU | Some sort of failure on the system at about 23:30. The fantech went from good values to bad values suggesting a fan failure. Also, while the measurements are close to 0, there was a dramatic change in the W/m2 value right at the same time as the fan change. THis could suggest a bigger failure. | "
IRT | generally follows the Vaisala T-2m very closely with only up to about 0.2 C difference. | "
SR50 | Distance getting larger? | "
PTU | All looks reasonable | "
Sonic | reasonable | "
Licor | NA | "
Flux Plates | reasonable | "
Camera_Rad | NR | Any pictures available?
Camera_EFOY | NR | Any pictures available?
Data Logger | max scans = 3317, Ptemp = -7 to +3C | Max scan time. Reasonable temperature range? Fans working?
Data Files
ASFS_30_slow | good | daily files of ~1120kB
ASFS_30_met | good | daily files of ~1740kB
ASFS_30_fast |  | daily files of ~5280kB
ASFS_30_Iridium_sumfile | NA | accumulating file with time stamp [include time stamp]

## ASFS50 (L3)
Atmospheric Surface Flux Stations in the Distributed Network
Note: Use this form for daily, remote assessments of ASFS. Use dedicated ASFS Site Visit form for site visits
Explanation
Communication System Status
Comms System | Iridium | [Radio, Iridium, or None] Indicate which forms of communication have been active
Quality | good enough | Provide any information on the quality of connection, how this impacts data flow
Power System Status
Operational Status | good | [Good, Questionable, Unknown]
Assessment of Exhaust | NR | Any notes about the buildup of ice that may be visible from the camera
Fuel Available | NR | Current remaining fuel available
Comments | 24V is good with similar behavior since the beginning. | Any other assessment of the power system, including quality of battery voltage/charge, operation of EFOY, etc.
Conditions
Structural | NR | Any useful information on the structural integrity of the ASFS
Position/Orientation/Level | heading steady, no major change in sonic inclination | Any useful information on position; i.e., rapid changes in GPS position, GPS orientation, or sonic-inclinometer level
Surface | NR | Any useful information on local surface conditions from camera
Sky-Weather | NR | Any useful information on weather or sky conditions
Instrument/Measurement Assessment
LWD | consistent with mostly clear day with brief cloudiness | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment from imagery if possible.
LWU | consistent with LWD | "
SWD | NA | "
SWU | NA | "
IRT | continued close tracking with Vaisala T | "
SR50 | clear temperature sensitivity in the SR50 reading | "
PTU | reasonable | "
Sonic | reasonable | "
Licor | NA | "
Flux Plates | Reasonable but two plates of opposite sign. | "
Camera_Rad | NR | Any pictures available?
Camera_EFOY | NR | Any pictures available?
Data Logger | Main counts=120, Ptemp=-3C | Max scan time. Reasonable temperature range? Fans working?
Data Files
ASFS_50_slow |  | daily files of ~1120kB
ASFS_50_met |  | daily files of ~1740kB
ASFS_50_fast |  | daily files of ~5280kB
ASFS_50_Iridium_sumfile |  | accumulating file with time stamp [include time stamp]

## ARM
ARM Facility - General Data Oversight
Explanation
Measurement Assessments | Assessment of data for scientific quality. Do the measurements look reasonable compared to conditions? Are the measurements capturing the key features?
Clouds
KAZR
MWACR
Ka-SACR
X-SACR
HSRL
MPL
Ceilometer
MWR2C
MWR3C
TSI
Precipitation
LDIS/PARS2 (ice)
LDIS/PARS2 (ship)
PWD (ice)
PWD (ship)
Siphon rain gauge (ship)
ORG (ice)
WBRG/Pluvio (ice)
Winds
BSRWP
DL
Radiation
AERI
ShipRad - PIR, PSP, SPN1
SkyRad - PSP, PIR, SPN1 (ice)
GndRad - PIR, PSP (ice)
MFR (ice)
CSPHOT
FRSR
Aerosol
AOSMet
CO / Purge
O3
CPC
CPCU
SMPS
HTDMA
UHSAS
CCNC
ACSM
SP2
PSAP
Nephelometer
INP filters
Data Requests | Any facilitation of MOSAiC data users

## Guests
Guest and Partner Instruments
Explanation
Power Distribution Unit - Met Hut (PFS)
Operational status | Operational and healthy.
Power balance on two PDUs? | Good balance, although still not much usage
Back-up generator test (/month) | NR
Precip VISSS - Met City (Maahn)
Backlight housings were clear of frost, which is an improvement over yesterday
Doppler Lidar - Met City (Brooks)
Doppler Lidar - Polarstern (Brooks)
Doppler Lidar - Polarstern (Heinemann)
Sodar - Met City (Brooks)
Radiation Suite - Met City (Pirazzini)
Not yet installed