## General
Surface Flux Team Daily Checks
If nothing to report use NR
Explanation
General
Date: | 20191101 | [YYYYMMDD]
Operator1: | Matt Shupe, Dave Costa | [name]
Operator2: | Ola Persson | [name]
Status of Operators: | NR | Any concerns or limitations regarding operators? Health, injury, etc.
General Weather Conditions: | light wind, making work around the tower easier. | Brief weather synopsis, especially weather that would impact operations
General Sea Ice Conditions: | monitoring cracks but not much change | Brief local sea ice synopsis, especially conditions that would impact operations
Special Operations: | NR | Any daily activities or special operations that would possibly impact daily schedule or operations
Command Center
Utilities: | Good | [comments or NR] Status on power connection, power panel, heaters, lights, network connection
Physical Integrity: | Good | [comments or NR] Any damage, ingress of water/snow, snow/water management, other structural concerns
DAQ Computer: | NR
Workstation Computer: | NR
Communications Mode(s): | Radio, Iridium, cable | [Radio, Iridium, Cable, and/or None] Which comms systems are being used?
Communications Quality: | Good generally, but lacking radio comms with L3 | Any information about operation of the radio and/or Iridium modems. What is the quality of connection? Any actions that were taken to modify the quality
Remote Mounts
Iridium Antenna - ARM Container | NR | Physical system assessment. Describe any actions taken such as cleaning or adjustments
Radio modem/antenna(s) - Crow's Nest | NR | Physical system assessment. Describe any actions taken such as cleaning or adjustments
Time-Lapse Camera - Crow's Nest | NA | Operational? Comment on visual assessment, and any actions taken. Data downloaded?
ACTION ITEMS
| VISSS calibration, finalize Finnish radiometer installation | List things that need to be accomplished in the coming days, developing issues, etc.
Supplies
| Schuko to US adapters, Tie wraps | Ongoing list of supplies needed in the field; carry this forward until supplies have been obtained

## Met Tower
Met Towers at Met City (11m tower and 30m mast)
Explanation
Local Conditions
Ice and Snow Conditions | light winds and not much change to ice conditions, monitoring prior cracks | Ice thickness, cracks, snow drifting, ablation or any other relevant conditions for operations.
Ice and Snow Management | NR | Anything required for physical maintenance of local area
Utility Status
Power line connectivity | up and robust | [up or down, comment on stability]
Network connectivity | good | [up or down, comment on stability]
Utility cables | good | [good or in need of maintenance]. Utility cables on poles or maintained routinely on surface, connections secure, strain relief intact, etc.
Control box | good | Any comments on the operations of the computer, data logger, power distribution, or other components. Temperature range and stability inside.
SPC Control box | NR | Any comments on Frey's power and control box.
Structural Assessment - 11m
Tower Base | Good | Stability of the base, including any potential ablation
Guy Anchors | Good | Any sign of anchor instability?
Guy Tightness | balanced | Even tightness? If not adjust.
Structural Integrity | Good | Any issues identified with structural components? Bolts, clamps, cables checked for tightness?
Plumbness | NR | Assess any shifts in plumbness of tower, possibly using manual measurements and/or sonic inclinometers
Boom Level / Stability | NR | Assess stability and level of booms
Cabling | Good | Secure with no rubbing.
Structural Assessment - 30m
Tower Base | NR | Stability of the base, including any potential ablation
Guy Anchors | NR | Any sign of anchor instability?
Guy Tightness | NR | Even tightness? If not adjust.
Structural Integrity | NR | Any issues identified with structural components? Bolts, clamps, cables checked for tightness?
Plumbness | NR | Assess any shifts in plumbness of tower, possibly using manual measurements and/or sonic inclinometers
Cabling | NR | Secure with no rubbing.
Instrument/Measurement Assessment
Daily Cleaning | Substantial tower activities today. Climbed half way up at about 05:00 UTC to clean rime/frost off of the upper parts of the tower, including all instruments and structure. Also moved the licor from 6m down to 2m height for re-installation. | Include the timing and any details related to cleaning
PTU_2m | 0435-0455 UTC removed ice buildup | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment, and any actions taken.
HMT_6m | removed ice buildup | "
HMT_11m | removed ice buildup | "
WXT_30m | NR | "
Sonic_2m | removed ice buildup from the base, but otherwise operating ice free | "
Sonic_6m | removed ice buildup from the base, but otherwise operating ice free | "
Sonic_11m | removed ice buildup from the base, but otherwise operating ice free | "
Sonic_30m | NR | "
Licor | At about 05:25 UTC moved this system from nominal 6m to 2m height and cleaned thoroughly with ethanol. Blinking light stopped. 0801 UTC - licor still looks clean | "
SR50 | Cleaned rime around the body, but sensing surface was clear | "
IRT | Cleaned rime around the body, but sensing surface was clear | "
Flux Plates | NR | "
CLASP | Cleaned rime off of instrument body and near inlet | "
SPCs (multiple) | Cleaned rime off of instrument arms and fins at both heights. | "
Data Files
Clasp02m | good | hourly files with size 3,095kB
CR1000X | good | daily file with size ~22500kB,
Licor02m | good | hourly files, raw with size ~3700-4000kB, stats with size ~19kB
Metek02m | good | hourly files, raw with size ~5280kB, stats with size 25kb
Metek06m | good | hourly files, raw with size ~5280kB, stats with size 25kb
Metek10m | good | hourly files, raw with size ~5280kB, stats with size 25kb
Metek30m | good | hourly files, raw with size ~5280kB, stats with size 25kb
Ozone | not operational | ????
Picarro | good | hourly files with size 3586kB

## ASFS40 (L1)
Atmospheric Surface Flux Stations in the Distributed Network
Note: Use this form for daily, remote assessments of ASFS. Use dedicated ASFS Site Visit form for site visits
Explanation
Communication System Status
Comms System | Radio | [Radio, Iridium, or None] Indicate which forms of communication have been active
Quality | good | Provide any information on the quality of connection, how this impacts data flow
Power System Status
Operational Status | Good | [Good, Questionable, Unknown]
Assessment of Exhaust | NR | Any notes about the buildup of ice that may be visible from the camera
Fuel Available | NR | Current remaining fuel available
Comments | No major change in EFOY power profile. Looks like 6 cycles over the day with the EFOY off for <1hr at a time. 24 and 12V look good | Any other assessment of the power system, including quality of battery voltage/charge, operation of EFOY, etc.
Conditions
Structural | NR | Any useful information on the structural integrity of the ASFS
Position/Orientation/Level | no major changes | Any useful information on position; i.e., rapid changes in GPS position, GPS orientation, or sonic-inclinometer level
Surface | NR | Any useful information on local surface conditions from camera
Sky-Weather | NR | Any useful information on weather or sky conditions
Instrument/Measurement Assessment
LWD | Consistent with clouds all day (similar to L3), fan good | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment from imagery if possible.
LWU | consistent with LWD, fan good | "
SWD | fan and heat good | "
SWU | fan and heat good | "
IRT | following but offset from T_2m by about 0.1C | "
SR50 | NR | "
PTU | reasonable | "
Sonic | reasonable | "
Licor | NA | "
Flux Plates | consistent and offset by 8 W/m2 | "
Camera_Rad | NR | Any pictures available?
Camera_EFOY | NR | Any pictures available?
Data Logger | mainscan<4000, Ptemp=-4 to 4C | Main scan time. Reasonable temperature range? Fans working?
Data Files
ASFS_40_slow | good | daily files of ~1120kB
ASFS_40_met | good | daily files of ~1740kB
ASFS_40_fast | good | daily files of ~100+MB
ASFS_40_Iridium_sumfile | not applicable | accumulating file with time stamp [include time stamp]

## ASFS30 (L2)
Atmospheric Surface Flux Stations in the Distributed Network
Note: Use this form for daily, remote assessments of ASFS. Use dedicated ASFS Site Visit form for site visits
Explanation
Communication System Status
Comms System | radio | [Radio, Iridium, or None] Indicate which forms of communication have been active
Quality | good | Provide any information on the quality of connection, how this impacts data flow
Power System Status
Operational Status | good | [Good, Questionable, Unknown]
Assessment of Exhaust | NR | Any notes about the buildup of ice that may be visible from the camera
Fuel Available | NR | Current remaining fuel available
Comments | EFOY power profile looks consistent with about two cycles/day and the system off for about 6 hours at a tmie. 24 and 12v good | Any other assessment of the power system, including quality of battery voltage/charge, operation of EFOY, etc.
Conditions
Structural | NR | Any useful information on the structural integrity of the ASFS
Position/Orientation/Level | no major change | Any useful information on position; i.e., rapid changes in GPS position, GPS orientation, or sonic-inclinometer level
Surface | NR | Any useful information on local surface conditions from camera
Sky-Weather | NR | Any useful information on weather or sky conditions
Instrument/Measurement Assessment
LWD | consistent with iced up radiometer, fan dead | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment from imagery if possible.
LWU | reasonable, fan good | "
SWD | fan and heat good | "
SWU | fan dead, heat good | "
IRT | about 0.1C warmer than T-2m | "
SR50 | NR | "
PTU | A couple of RH spikes early in the day that may be suspect for exhaust contamination | "
Sonic | reasonable | "
Licor | NA | "
Flux Plates | opposite signs and very similar magnitudes | "
Camera_Rad | NR | Any pictures available?
Camera_EFOY | NR | Any pictures available?
Data Logger | mainscan < 3400, Ptemp = -5 to 4C | Main scan time. Reasonable temperature range? Fans working?
Data Files
ASFS_30_slow | good | daily files of ~1120kB
ASFS_30_met | good | daily files of ~1740kB
ASFS_30_fast | good | daily files of ~5280kB
ASFS_30_Iridium_sumfile | no iridium connection | accumulating file with time stamp [include time stamp]

## ASFS50 (L3)
Atmospheric Surface Flux Stations in the Distributed Network
Note: Use this form for daily, remote assessments of ASFS. Use dedicated ASFS Site Visit form for site visits
Explanation
Communication System Status
Comms System | Iridium | [Radio, Iridium, or None] Indicate which forms of communication have been active
Quality | good | Provide any information on the quality of connection, how this impacts data flow
Power System Status
Operational Status | good | [Good, Questionable, Unknown]
Assessment of Exhaust | NR | Any notes about the buildup of ice that may be visible from the camera
Fuel Available | NR | Current remaining fuel available
Comments | limited EFOY data making it in, but the available data suggests that the 24V is doing ok. | Any other assessment of the power system, including quality of battery voltage/charge, operation of EFOY, etc.
Conditions
Structural | NR | Any useful information on the structural integrity of the ASFS
Position/Orientation/Level | inclinometer is -0.07 and -0.43deg, | Any useful information on position; i.e., rapid changes in GPS position, GPS orientation, or sonic-inclinometer level
Surface | NR | Any useful information on local surface conditions from camera
Sky-Weather | NR | Any useful information on weather or sky conditions
Instrument/Measurement Assessment
LWD | Cloudy all day or icing on radiometer, fan good | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment from imagery if possible.
LWU | consistent with LWD, fan good | "
SWD | heat good. Fantach had a sharp decrease | "
SWU | heat and fan good | "
IRT | follows T_2m with small offset | "
SR50 | NR | "
PTU | reasonable | "
Sonic | reasonable | "
Licor | NA | "
Flux Plates | NR | "
Camera_Rad | NR | Any pictures available?
Camera_EFOY | NR | Any pictures available?
Data Logger | Ptemp=-4 to 0C, main counts=120 | Main scan time. Reasonable temperature range? Fans working?
Data Files
ASFS_50_slow | no radio | daily files of ~1120kB
ASFS_50_met | no radio | daily files of ~1740kB
ASFS_50_fast | no radio | daily files of ~5280kB
ASFS_50_Iridium_sumfile | good | accumulating file with time stamp [include time stamp]

## ARM
ARM Facility - General Data Oversight
Explanation
Measurement Assessments | Assessment of data for scientific quality. Do the measurements look reasonable compared to conditions? Are the measurements capturing the key features?
Clouds
KAZR
MWACR
Ka-SACR
X-SACR
HSRL
MPL
Ceilometer
MWR2C
MWR3C
TSI
Precipitation
LDIS/PARS2 (ice)
LDIS/PARS2 (ship)
PWD (ice)
PWD (ship)
Siphon rain gauge (ship)
ORG (ice)
WBRG/Pluvio (ice)
Winds
BSRWP
DL
Radiation
AERI
ShipRad - PIR, PSP, SPN1
SkyRad - PSP, PIR, SPN1 (ice)
GndRad - PIR, PSP (ice)
MFR (ice)
CSPHOT
FRSR
Aerosol
AOSMet
CO / Purge
O3
CPC
CPCU
SMPS
HTDMA
UHSAS
CCNC
ACSM
SP2
PSAP
Nephelometer
INP filters
Data Requests | Any facilitation of MOSAiC data users

## Guests
Guest and Partner Instruments
Explanation
Power Distribution Unit - Met Hut (PFS)
Operational status | operational
Power balance on two PDUs? | NR
Back-up generator test (/month) | not done
Precip VISSS - Met City (Maahn)
Cleaned backlights at about 08:15 UTC
Doppler Lidar - Met City (Brooks)
Brooks onsite
Doppler Lidar - Polarstern (Brooks)
Brooks onsite
Doppler Lidar - Polarstern (Heinemann)
Brooks onsite
Sodar - Met City (Brooks)
Brooks onsite
Radiation Suite - Met City (Pirazzini)
Set up the radiation suite on the ice at about 07:00 UTC, including the stakes into the ice with the base plate, the stakes supporting the guy lines, and the basic guy line tensioning to keep the system plumb and level.