## General
Surface Flux Team Daily Checks
If nothing to report use NR
Explanation
General
Date: | 20191102 | [YYYYMMDD]
Operator1: | Matt Shupe, Dave Costa | [name]
Operator2: | Ola Persson | [name]
Status of Operators: | NR | Any concerns or limitations regarding operators? Health, injury, etc.
General Weather Conditions: | Very little wind, making for a beautiful and comfortable day. | Brief weather synopsis, especially weather that would impact operations
General Sea Ice Conditions: | Cracks through camp opened again. 4-5cm under the ARM swingset with hairline crack under the VISSS and near the met tower cables. Monitored this at multiple times during the day | Brief local sea ice synopsis, especially conditions that would impact operations
Special Operations: | Met City visit 01:00-02:30 UTC, quick Met City visit for crack monitoring at about 07:30 UTC | Any daily activities or special operations that would possibly impact daily schedule or operations
Command Center
Utilities: | NR | [comments or NR] Status on power connection, power panel, heaters, lights, network connection
Physical Integrity: | NR | [comments or NR] Any damage, ingress of water/snow, snow/water management, other structural concerns
DAQ Computer: | NR
Workstation Computer: | NR
Communications Mode(s): | NR | [Radio, Iridium, Cable, and/or None] Which comms systems are being used?
Communications Quality: | NR | Any information about operation of the radio and/or Iridium modems. What is the quality of connection? Any actions that were taken to modify the quality
Remote Mounts
Iridium Antenna - ARM Container | NR | Physical system assessment. Describe any actions taken such as cleaning or adjustments
Radio modem/antenna(s) - Crow's Nest | NR | Physical system assessment. Describe any actions taken such as cleaning or adjustments
Time-Lapse Camera - Crow's Nest | NR | Operational? Comment on visual assessment, and any actions taken. Data downloaded?
ACTION ITEMS
| VISSS calibration, Finnish radiation calibration, L2 radiometer preparation | List things that need to be accomplished in the coming days, developing issues, etc.
Supplies
| Schuko to US adapters, Tie wraps | Ongoing list of supplies needed in the field; carry this forward until supplies have been obtained

## Met Tower
Met Towers at Met City (11m tower and 30m mast)
Explanation
Local Conditions
Ice and Snow Conditions | Monitoring small cracks around met city. The crack to the south of the tower openned again, up to about 8 cm but moving slowly and then stopped. Some snow buildup around camp with small drifts; photos taken of cracks | Ice thickness, cracks, snow drifting, ablation or any other relevant conditions for operations.
Ice and Snow Management | Removed some drifting around the base of the tower | Anything required for physical maintenance of local area
Utility Status
Power line connectivity | Up | [up or down, comment on stability]
Network connectivity | Up | [up or down, comment on stability]
Utility cables | Good | [good or in need of maintenance]. Utility cables on poles or maintained routinely on surface, connections secure, strain relief intact, etc.
Control box | Good | Any comments on the operations of the computer, data logger, power distribution, or other components. Temperature range and stability inside.
SPC Control box | Good | Any comments on Frey's power and control box.
Structural Assessment - 11m
Tower Base | NR | Stability of the base, including any potential ablation
Guy Anchors | NR | Any sign of anchor instability?
Guy Tightness | NR | Even tightness? If not adjust.
Structural Integrity | NR | Any issues identified with structural components? Bolts, clamps, cables checked for tightness?
Plumbness | NR | Assess any shifts in plumbness of tower, possibly using manual measurements and/or sonic inclinometers
Boom Level / Stability | NR | Assess stability and level of booms
Cabling | moved some of the cabling along the power poes to support slack over the small crack through camp | Secure with no rubbing.
Structural Assessment - 30m
Tower Base | NR | Stability of the base, including any potential ablation
Guy Anchors | NR | Any sign of anchor instability?
Guy Tightness | NR | Even tightness? If not adjust.
Structural Integrity | NR | Any issues identified with structural components? Bolts, clamps, cables checked for tightness?
Plumbness | NR | Assess any shifts in plumbness of tower, possibly using manual measurements and/or sonic inclinometers
Cabling | moved some of the cabling along the ground to support slack over the small crack | Secure with no rubbing.
Instrument/Measurement Assessment
Daily Cleaning | Tower cleaning at 00:30-00:40 UTC. There was little to no icing. Took the opportunity to clean off other prior icing from tower structure and add more slack to cables | Include the timing and any details related to cleaning
PTU_2m | good | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment, and any actions taken.
HMT_6m | good | "
HMT_11m | good | "
WXT_30m | NR | "
Sonic_2m | clean | "
Sonic_6m | clean | "
Sonic_11m | clean | "
Sonic_30m | Looks like the Metek data stopped at about 00:35 UTC, which was right during the tower cleaning and maintenance time period. This could be related to a connector coming un-fastened, possibly during cable maintenance (i.e., re-securing and adjusting cables on power poles to make slack). | "
Licor | 0041 UTC, For the first time in many days (ever?), the licor light was NOT blinking, suggesting that there were no errors. The system looked to be free of icing. But still cleaned both lenses with ethanol | "
SR50 | clean | "
IRT | clean | "
Flux Plates | areas around flux plates have been evolving somewhat as there is more snow and drifting and then some scouring. | "
CLASP | clean | "
SPCs (multiple) | clean | "
Data Files
Clasp02m | good | hourly files with size 3,095kB
CR1000X | good | daily file with size ~22500kB,
Licor02m | good | hourly files, raw with size ~3700-4000kB, stats with size ~19kB
Metek02m | good | hourly files, raw with size ~5280kB, stats with size 25kb
Metek06m | good | hourly files, raw with size ~5280kB, stats with size 25kb
Metek10m | good | hourly files, raw with size ~5280kB, stats with size 25kb
Metek30m | No data after 00:36 UTC | hourly files, raw with size ~5280kB, stats with size 25kb
Ozone | not operational | ????
Picarro | good | hourly files with size 3586kB

## ASFS40 (L1)
Atmospheric Surface Flux Stations in the Distributed Network
Note: Use this form for daily, remote assessments of ASFS. Use dedicated ASFS Site Visit form for site visits
Explanation
Communication System Status
Comms System | radio | [Radio, Iridium, or None] Indicate which forms of communication have been active
Quality | good | Provide any information on the quality of connection, how this impacts data flow
Power System Status
Operational Status | good | [Good, Questionable, Unknown]
Assessment of Exhaust | NR | Any notes about the buildup of ice that may be visible from the camera
Fuel Available | NR | Current remaining fuel available
Comments | EFOY power profile looks normal, 24 and 12 V good | Any other assessment of the power system, including quality of battery voltage/charge, operation of EFOY, etc.
Conditions
Structural | NR | Any useful information on the structural integrity of the ASFS
Position/Orientation/Level | no major changes | Any useful information on position; i.e., rapid changes in GPS position, GPS orientation, or sonic-inclinometer level
Surface | NR | Any useful information on local surface conditions from camera
Sky-Weather | NR | Any useful information on weather or sky conditions
Instrument/Measurement Assessment
LWD | mostly cloudy with periodic clearing. fan good | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment from imagery if possible.
LWU | reasonable, fan good | "
SWD | fan and heat good | "
SWU | fan and heat good | "
IRT | follow T_2m but about 0.1-0.4C colder | "
SR50 | NR | "
PTU | good | "
Sonic | reasonable | "
Licor | NA | "
Flux Plates | similar signal with ofset of about 7 W/m2 | "
Camera_Rad | NR | Any pictures available?
Camera_EFOY | NR | Any pictures available?
Data Logger | mainscan = <4000, Ptemp=-1 to 7 C | Main scan time. Reasonable temperature range? Fans working?
Data Files
ASFS_40_slow | good | daily files of ~1120kB
ASFS_40_met | good | daily files of ~1740kB
ASFS_40_fast | only 6 hrs of data | daily files of ~5280kB
ASFS_40_Iridium_sumfile | no iridium | accumulating file with time stamp [include time stamp]

## ASFS30 (L2)
Atmospheric Surface Flux Stations in the Distributed Network
Note: Use this form for daily, remote assessments of ASFS. Use dedicated ASFS Site Visit form for site visits
Explanation
Communication System Status
Comms System | radio | [Radio, Iridium, or None] Indicate which forms of communication have been active
Quality | generally good | Provide any information on the quality of connection, how this impacts data flow
Power System Status
Operational Status | good, except for failures outlined below | [Good, Questionable, Unknown]
Assessment of Exhaust | NR | Any notes about the buildup of ice that may be visible from the camera
Fuel Available | NR | Current remaining fuel available
Comments | EFOY power profile looks consistent, 24 and 12 V look good | Any other assessment of the power system, including quality of battery voltage/charge, operation of EFOY, etc.
Conditions
Structural | NR | Any useful information on the structural integrity of the ASFS
Position/Orientation/Level | no major changes | Any useful information on position; i.e., rapid changes in GPS position, GPS orientation, or sonic-inclinometer level
Surface | NR | Any useful information on local surface conditions from camera
Sky-Weather | NR | Any useful information on weather or sky conditions
Instrument/Measurement Assessment
LWD | iced over, fan dead | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment from imagery if possible.
LWU | reasonable measurements, fan good | "
SWD | heat and fan good | "
SWU | heat good, fan dead | "
IRT | IRT tracks T_2m but is 0.2-0.5C warmer | "
SR50 | NR | "
PTU | some large swings in RH | "
Sonic | reasonable measurements, | "
Licor | NA | "
Flux Plates | opposite signs and similar magnitudes | "
Camera_Rad | NR | Any pictures available?
Camera_EFOY | NR | Any pictures available?
Data Logger | mainscans = <3400, Ptemp=-3 to 3 C | Main scan time. Reasonable temperature range? Fans working?
Data Files
ASFS_30_slow | good | daily files of ~1120kB
ASFS_30_met | good | daily files of ~1740kB
ASFS_30_fast | only ~6hrs of data | daily files of ~5280kB
ASFS_30_Iridium_sumfile | no iridium | accumulating file with time stamp [include time stamp]

## ASFS50 (L3)
Atmospheric Surface Flux Stations in the Distributed Network
Note: Use this form for daily, remote assessments of ASFS. Use dedicated ASFS Site Visit form for site visits
Explanation
Communication System Status
Comms System | Iridium | [Radio, Iridium, or None] Indicate which forms of communication have been active
Quality | good | Provide any information on the quality of connection, how this impacts data flow
Power System Status
Operational Status | good | [Good, Questionable, Unknown]
Assessment of Exhaust | NR | Any notes about the buildup of ice that may be visible from the camera
Fuel Available | NR | Current remaining fuel available
Comments | limited EFOY data making it in, but the available data suggests that the 24V is doing ok. | Any other assessment of the power system, including quality of battery voltage/charge, operation of EFOY, etc.
Conditions
Structural | NR | Any useful information on the structural integrity of the ASFS
Position/Orientation/Level | inclination=-0.07 and -0.43deg, | Any useful information on position; i.e., rapid changes in GPS position, GPS orientation, or sonic-inclinometer level
Surface | NR | Any useful information on local surface conditions from camera
Sky-Weather | NR | Any useful information on weather or sky conditions
Instrument/Measurement Assessment
LWD | Cloudy or iced up, fan good | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment from imagery if possible.
LWU | fan good | "
SWD | heat good, fan operating at a lower fantach rate | "
SWU | heat and fan good | "
IRT | follows T_2m with small offset | "
SR50 | NR | "
PTU | reasonable | "
Sonic | reasonable | "
Licor | NA | "
Flux Plates | NR | "
Camera_Rad | NR | Any pictures available?
Camera_EFOY | NR | Any pictures available?
Data Logger | Ptemp=-2 to 0 C, main counts=120 | Main scan time. Reasonable temperature range? Fans working?
Data Files
ASFS_50_slow | no radio | daily files of ~1120kB
ASFS_50_met | no radio | daily files of ~1740kB
ASFS_50_fast | no radio | daily files of ~5280kB
ASFS_50_Iridium_sumfile | only 6 hours of data, loggernet shut off for a while | accumulating file with time stamp [include time stamp]

## ARM
ARM Facility - General Data Oversight
swing set anchors on NW side ripped out by small crack under pad
Explanation
Measurement Assessments |  | Assessment of data for scientific quality. Do the measurements look reasonable compared to conditions? Are the measurements capturing the key features?
Clouds
KAZR
MWACR
Ka-SACR
X-SACR
HSRL
MPL
Ceilometer
MWR2C
MWR3C
TSI
Precipitation
LDIS/PARS2 (ice)
LDIS/PARS2 (ship)
PWD (ice)
PWD (ship)
Siphon rain gauge (ship)
ORG (ice)
WBRG/Pluvio (ice)
Winds
BSRWP
DL
Radiation
AERI
ShipRad - PIR, PSP, SPN1
SkyRad - PSP, PIR, SPN1 (ice)
GndRad - PIR, PSP (ice)
MFR (ice)
CSPHOT
FRSR
Aerosol
AOSMet
CO / Purge
O3
CPC
CPCU
SMPS
HTDMA
UHSAS
CCNC
ACSM
SP2
PSAP
Nephelometer
INP filters
Data Requests |  | Any facilitation of MOSAiC data users

## Guests
Guest and Partner Instruments
Explanation
Power Distribution Unit - Met Hut (PFS)
Operational status | Good
Power balance on two PDUs? | Good
Back-up generator test (/month) | Not done yet
Precip VISSS - Met City (Maahn)
Small crack (<1cm) near one of the guy lines. We moved this to the other side of the crack to minimize potential tension.
Doppler Lidar - Met City (Brooks)
Brooks onsite
Doppler Lidar - Polarstern (Brooks)
Brooks onsite
Doppler Lidar - Polarstern (Heinemann)
Brooks onsite
Sodar - Met City (Brooks)
Brooks onsite
Radiation Suite - Met City (Pirazzini)
Started up the system at 01:50 UTC, levelled at 02:24 UTC