## General
Surface Flux Team Daily Checks
If nothing to report use NR
Explanation
General
Date: | 20191105 | [YYYYMMDD]
Operator1: | Matt Shupe, Dave Costa | [name]
Operator2: | Ola Persson | [name]
Status of Operators: | NR | Any concerns or limitations regarding operators? Health, injury, etc.
General Weather Conditions: | mostly clear and cold | Brief weather synopsis, especially weather that would impact operations
General Sea Ice Conditions: | unchanged | Brief local sea ice synopsis, especially conditions that would impact operations
Special Operations: | Met City visit today at 01:00-02:30 UTC | Any daily activities or special operations that would possibly impact daily schedule or operations
Command Center
Utilities: | All good | [comments or NR] Status on power connection, power panel, heaters, lights, network connection
Physical Integrity: | NR | [comments or NR] Any damage, ingress of water/snow, snow/water management, other structural concerns
DAQ Computer: | NR
Workstation Computer: | NR
Communications Mode(s): | Radio, Iridium, Cable | [Radio, Iridium, Cable, and/or None] Which comms systems are being used?
Communications Quality: | Ok to the ice, L3 is on good Iridium comms, Declined radio connection quality to L2. L1 is still solid | Any information about operation of the radio and/or Iridium modems. What is the quality of connection? Any actions that were taken to modify the quality
Remote Mounts
Iridium Antenna - ARM Container | Looks clean. | Physical system assessment. Describe any actions taken such as cleaning or adjustments
Radio modem/antenna(s) - Crow's Nest | Cleaned a thick layer of ice (~1") off of the starboard antenna and most of a thin layer of rime off of the port side antenna. Measured antenna lengths: 34" wth 1.5" diameter | Physical system assessment. Describe any actions taken such as cleaning or adjustments
Time-Lapse Camera - Crow's Nest | NA | Operational? Comment on visual assessment, and any actions taken. Data downloaded?
ACTION ITEMS
| VISSS calibration, FMI radiometer suite calibration, climb 11m tower for icing mitigation on upper instruments and to get a current measure of all instrument heights. | List things that need to be accomplished in the coming days, developing issues, etc.
Supplies
| Schuko to US adapters, Tie wraps | Ongoing list of supplies needed in the field; carry this forward until supplies have been obtained

## Met Tower
Met Towers at Met City (11m tower and 30m mast)
Explanation
Local Conditions
Ice and Snow Conditions | Mild drifting with little other change | Ice thickness, cracks, snow drifting, ablation or any other relevant conditions for operations.
Ice and Snow Management | No action | Anything required for physical maintenance of local area
Utility Status
Power line connectivity | up and stable | [up or down, comment on stability]
Network connectivity | up and stable | [up or down, comment on stability]
Utility cables | good | [good or in need of maintenance]. Utility cables on poles or maintained routinely on surface, connections secure, strain relief intact, etc.
Control box | good | Any comments on the operations of the computer, data logger, power distribution, or other components. Temperature range and stability inside.
SPC Control box | NR | Any comments on Frey's power and control box.
Structural Assessment - 11m
Tower Base | NR | Stability of the base, including any potential ablation
Guy Anchors | NR | Any sign of anchor instability?
Guy Tightness | NR | Even tightness? If not adjust.
Structural Integrity | No visible issues | Any issues identified with structural components? Bolts, clamps, cables checked for tightness?
Plumbness | NR | Assess any shifts in plumbness of tower, possibly using manual measurements and/or sonic inclinometers
Boom Level / Stability | NR | Assess stability and level of booms
Cabling | NR | Secure with no rubbing.
Structural Assessment - 30m
Tower Base | NR | Stability of the base, including any potential ablation
Guy Anchors | NR | Any sign of anchor instability?
Guy Tightness | NR | Even tightness? If not adjust.
Structural Integrity | No visible issues | Any issues identified with structural components? Bolts, clamps, cables checked for tightness?
Plumbness | NR | Assess any shifts in plumbness of tower, possibly using manual measurements and/or sonic inclinometers
Cabling | NR | Secure with no rubbing.
Instrument/Measurement Assessment
Daily Cleaning | Site visit 01:00-02:30. Cleaning instruments at about 01:30 | Include the timing and any details related to cleaning
PTU_2m | icing cleared | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment, and any actions taken.
HMT_6m | icing cleared | "
HMT_11m | totally iced over | "
WXT_30m | NR | "
Sonic_2m | no icing | "
Sonic_6m | no icing | "
Sonic_11m | no icing | "
Sonic_30m | no icing, found a loose connection and re-connected and secured. Data started again. | "
Licor | 0137 UTC cleaned Licor (some rime crystals - took photo); 195 cm from base of tower to licor center; 35 cm licor center to Metek center | "
SR50 | Free of icing | "
IRT | Free of icing | "
Flux Plates | NR | "
CLASP | Free of icing | "
SPCs (multiple) | Free of icing | "
Data Files
Clasp02m | good | hourly files with size 3,095kB
CR1000X | good | daily file with size ~22500kB,
Licor02m | good | hourly files, raw with size ~3700-4000kB, stats with size ~19kB
Metek02m | good | hourly files, raw with size ~5280kB, stats with size 25kb
Metek06m | good | hourly files, raw with size ~5280kB, stats with size 25kb
Metek10m | good | hourly files, raw with size ~5280kB, stats with size 25kb
Metek30m | data started after 01:00 UTC | hourly files, raw with size ~5280kB, stats with size 25kb
Ozone | NA | ????
Picarro | good | hourly files with size 3586kB

## ASFS40 (L1)
Atmospheric Surface Flux Stations in the Distributed Network
Note: Use this form for daily, remote assessments of ASFS. Use dedicated ASFS Site Visit form for site visits
Explanation
Communication System Status
Comms System | radio | [Radio, Iridium, or None] Indicate which forms of communication have been active
Quality | good | Provide any information on the quality of connection, how this impacts data flow
Power System Status
Operational Status | good | [Good, Questionable, Unknown]
Assessment of Exhaust | NR | Any notes about the buildup of ice that may be visible from the camera
Fuel Available | NR | Current remaining fuel available
Comments | EFOY power profile looks like there was some missing data in the middle of the day? But otherwise 24 and 12V are good. | Any other assessment of the power system, including quality of battery voltage/charge, operation of EFOY, etc.
Conditions
Structural | NR | Any useful information on the structural integrity of the ASFS
Position/Orientation/Level | no major changes | Any useful information on position; i.e., rapid changes in GPS position, GPS orientation, or sonic-inclinometer level
Surface | NR | Any useful information on local surface conditions from camera
Sky-Weather | NR | Any useful information on weather or sky conditions
Instrument/Measurement Assessment
LWD | Mostly clear, with clouds later in day. Fan good, with strange oscillation during the day | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment from imagery if possible.
LWU | Consistent with LWD. Fan good, with strange oscillation during the day | "
SWD | fan and heat good with strange oscillatoin during the day | "
SWU | fan and heat good with strange oscillatoin during the day | "
IRT | consistent with T_2m but generally colder by about 0.2-0.3C | "
SR50 | NR | "
PTU | reasonable | "
Sonic | reasonable | "
Licor | NA | "
Flux Plates | following each other with offset of up to 10 W/m2 | "
Camera_Rad | NR | Any pictures available?
Camera_EFOY | NR | Any pictures available?
Data Logger | mainscan < 4000, Ptemp=-7 to 2C | Main scan time. Reasonable temperature range? Fans working?
Data Files
ASFS_40_slow | good | daily files of ~1120kB
ASFS_40_met | good | daily files of ~1740kB
ASFS_40_fast | good | daily files of >105MB
ASFS_40_Iridium_sumfile | no iridium | accumulating file with time stamp [include time stamp]

## ASFS30 (L2)
Atmospheric Surface Flux Stations in the Distributed Network
Note: Use this form for daily, remote assessments of ASFS. Use dedicated ASFS Site Visit form for site visits
Explanation
Communication System Status
Comms System | radio | [Radio, Iridium, or None] Indicate which forms of communication have been active
Quality | ok,but getting worse | Provide any information on the quality of connection, how this impacts data flow
Power System Status
Operational Status | good | [Good, Questionable, Unknown]
Assessment of Exhaust | NR | Any notes about the buildup of ice that may be visible from the camera
Fuel Available | NR | Current remaining fuel available
Comments | EFOY power profile normal at almost 2 cycles per day. 24 and 12V good. | Any other assessment of the power system, including quality of battery voltage/charge, operation of EFOY, etc.
Conditions
Structural | NR | Any useful information on the structural integrity of the ASFS
Position/Orientation/Level | No big changes. Y inclination is -1.55deg. | Any useful information on position; i.e., rapid changes in GPS position, GPS orientation, or sonic-inclinometer level
Surface | NR | Any useful information on local surface conditions from camera
Sky-Weather | NR | Any useful information on weather or sky conditions
Instrument/Measurement Assessment
LWD | iced over, fan bad | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment from imagery if possible.
LWU | looks good, fan good | "
SWD | fan and heat good | "
SWU | fan bad, heat good | "
IRT | follows T-2m within 0.1 C, slightly warmer | "
SR50 | NR | "
PTU | RH jump at about 04:30 seems strange, | "
Sonic | reasonable | "
Licor | NA | "
Flux Plates | Opposite and offset of about 2 W/m2 | "
Camera_Rad | NR | Any pictures available?
Camera_EFOY | NR | Any pictures available?
Data Logger | mainscans= <4000, Ptemp=-10 to 0C | Main scan time. Reasonable temperature range? Fans working?
Data Files
ASFS_30_slow | good | daily files of ~1120kB
ASFS_30_met | good | daily files of ~1740kB
ASFS_30_fast | good | daily files of ~5280kB
ASFS_30_Iridium_sumfile | not available | accumulating file with time stamp [include time stamp]

## ASFS50 (L3)
Atmospheric Surface Flux Stations in the Distributed Network
Note: Use this form for daily, remote assessments of ASFS. Use dedicated ASFS Site Visit form for site visits
Explanation
Communication System Status
Comms System | iridium | [Radio, Iridium, or None] Indicate which forms of communication have been active
Quality | good | Provide any information on the quality of connection, how this impacts data flow
Power System Status
Operational Status | Good | [Good, Questionable, Unknown]
Assessment of Exhaust | NR | Any notes about the buildup of ice that may be visible from the camera
Fuel Available | NR | Current remaining fuel available
Comments | Not getting good power information presently. | Any other assessment of the power system, including quality of battery voltage/charge, operation of EFOY, etc.
Conditions
Structural | NR | Any useful information on the structural integrity of the ASFS
Position/Orientation/Level | no major change | Any useful information on position; i.e., rapid changes in GPS position, GPS orientation, or sonic-inclinometer level
Surface | NR | Any useful information on local surface conditions from camera
Sky-Weather | NR | Any useful information on weather or sky conditions
Instrument/Measurement Assessment
LWD | mostly clear with periodic clouds. fan good | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment from imagery if possible.
LWU | consistent with LWD. fan good | "
SWD | heat and fan good, oscillations in fantach have stopped. | "
SWU | heat and fan good | "
IRT | follows T_2m and is about 0.1C warmer. Not the same delay as in the other stations. | "
SR50 | surface continues to get farther away as temperatures get colder! | "
PTU | reasonable with some potential jumps | "
Sonic | reasonable with some potential jumps | "
Licor | NA | "
Flux Plates | opposite and different by about 1 W/m2 | "
Camera_Rad | NR | Any pictures available?
Camera_EFOY | NR | Any pictures available?
Data Logger | main counts = 120, Ptemp=-7 to -4C | Main scan time. Reasonable temperature range? Fans working?
Data Files
ASFS_50_slow | no radio | daily files of ~1120kB
ASFS_50_met | no radio | daily files of ~1740kB
ASFS_50_fast | no radio | daily files of ~5280kB
ASFS_50_Iridium_sumfile | good | accumulating file with time stamp [include time stamp]

## ARM
ARM Facility - General Data Oversight
Explanation
Measurement Assessments |  | Assessment of data for scientific quality. Do the measurements look reasonable compared to conditions? Are the measurements capturing the key features?
Clouds
KAZR
MWACR
Ka-SACR
X-SACR
HSRL
MPL
Ceilometer
MWR2C
MWR3C
TSI
Precipitation
LDIS/PARS2 (ice)
LDIS/PARS2 (ship)
PWD (ice)
PWD (ship)
Siphon rain gauge (ship)
ORG (ice)
WBRG/Pluvio (ice)
Winds
BSRWP
DL
Radiation
AERI
ShipRad - PIR, PSP, SPN1
SkyRad - PSP, PIR, SPN1 (ice)
GndRad - PIR, PSP (ice) | Measured the ice thickness near the swingset. For the N legs the ice was 154cm thick, near the S legs the ice was 123cm thick.
MFR (ice)
CSPHOT
FRSR
Aerosol
AOSMet
CO / Purge
O3
CPC
CPCU
SMPS
HTDMA
UHSAS
CCNC
ACSM
SP2
PSAP
Nephelometer
INP filters
Data Requests |  | Any facilitation of MOSAiC data users

## Guests
Guest and Partner Instruments
Explanation
Power Distribution Unit - Met Hut (PFS)
Operational status | Functional without alarms
Power balance on two PDUs? | NR
Back-up generator test (/month) | Not done
Precip VISSS - Met City (Maahn)
moderate icing buildup on the backlights, software all still functional
Doppler Lidar - Met City (Brooks)
Brooks onsite
Doppler Lidar - Polarstern (Brooks)
Brooks onsite
Doppler Lidar - Polarstern (Heinemann)
Brooks onsite
Sodar - Met City (Brooks)
Brooks onsite
Radiation Suite - Met City (Pirazzini)
All up looking radiometer domes again frosted over. Cleaned only the PIR uplooking at about 02:10UTC using a little ethanol and kimwipe. Put the battery power unit / sled into a good bag to mitigate snow build up; this is preferable to the tarp provided.