## General
Surface Flux Team Daily Checks
If nothing to report use NR
Explanation
General
Date: | 20191121 | [YYYYMMDD]
Operator1: | Matt Shupe, Dave Costa | [name]
Operator2: | Ola Persson | [name]
Status of Operators: | NR | Any concerns or limitations regarding operators? Health, injury, etc.
General Weather Conditions: | some low clouds, mild winds | Brief weather synopsis, especially weather that would impact operations
General Sea Ice Conditions: | stable today after many days with a lot of dynamics. | Brief local sea ice synopsis, especially conditions that would impact operations
Special Operations: | Visits to Met City today: 03:30-05:30 and 10:00-11:30 UTC. Helicopter operations for visiting L2 are on hold. | Any daily activities or special operations that would possibly impact daily schedule or operations
Command Center
Utilities: | good | [comments or NR] Status on power connection, power panel, heaters, lights, network connection
Physical Integrity: | good | [comments or NR] Any damage, ingress of water/snow, snow/water management, other structural concerns
DAQ Computer: | good
Workstation Computer: | good
Communications Mode(s): | radio, iridium. Cable connection to met City is down. | [Radio, Iridium, Cable, and/or None] Which comms systems are being used?
Communications Quality: | Good for the systems that are operational | Any information about operation of the radio and/or Iridium modems. What is the quality of connection? Any actions that were taken to modify the quality
Remote Mounts
Iridium Antenna - ARM Container | NR | Physical system assessment. Describe any actions taken such as cleaning or adjustments
Radio modem/antenna(s) - Crow's Nest | NR | Physical system assessment. Describe any actions taken such as cleaning or adjustments
Time-Lapse Camera - Crow's Nest | NA | Operational? Comment on visual assessment, and any actions taken. Data downloaded?
ACTION ITEMS
List things that need to be accomplished in the coming days, developing issues, etc.
Supplies
Ongoing list of supplies needed in the field; carry this forward until supplies have been obtained

## Met Tower
Met Towers at Met City (11m tower and 30m mast)
Explanation
Local Conditions
Ice and Snow Conditions | Stable ice conditions around Met City. The crack along the spine road widened to about 2-3m. This will require a new road be made to Met City. | Ice thickness, cracks, snow drifting, ablation or any other relevant conditions for operations.
Ice and Snow Management | A small amount of drift management around the Met Hut. | Anything required for physical maintenance of local area
Utility Status
Power line connectivity | down. Running on generators starting at about 03:45 throughout most of the rest of the day. | [up or down, comment on stability]
Network connectivity | down. | [up or down, comment on stability]
Utility cables | local cables are good. | [good or in need of maintenance]. Utility cables on poles or maintained routinely on surface, connections secure, strain relief intact, etc.
Control box | good | Any comments on the operations of the computer, data logger, power distribution, or other components. Temperature range and stability inside.
SPC Control box | NR | Any comments on Frey's power and control box.
Structural Assessment - 11m
Tower Base | stable | Stability of the base, including any potential ablation
Guy Anchors | stable. The guy anchors that were in danger from the crack appear to no longer be in danger as the crack has not moved in a number of days. We will see how things go in the coming days and then possibly reconnect to the better anchors. | Any sign of anchor instability?
Guy Tightness | even | Even tightness? If not adjust.
Structural Integrity | good | Any issues identified with structural components? Bolts, clamps, cables checked for tightness?
Plumbness | NR | Assess any shifts in plumbness of tower, possibly using manual measurements and/or sonic inclinometers
Boom Level / Stability | NR | Assess stability and level of booms
Cabling | good | Secure with no rubbing.
Structural Assessment - 30m
Tower Base | not installed | Stability of the base, including any potential ablation
Guy Anchors | NA | Any sign of anchor instability?
Guy Tightness | NA | Even tightness? If not adjust.
Structural Integrity | NA | Any issues identified with structural components? Bolts, clamps, cables checked for tightness?
Plumbness | NA | Assess any shifts in plumbness of tower, possibly using manual measurements and/or sonic inclinometers
Cabling | NA | Secure with no rubbing.
Instrument/Measurement Assessment
Daily Cleaning | Onsite at about 03:30 UTC with tower cleaning closer to 04:30 UTC. Most equipment has some icing as the system was not running for a number of hours prior to our arrival. | Include the timing and any details related to cleaning
PTU_2m | removed a little icing | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment, and any actions taken.
HMT_6m | appears to be some icing | "
HMT_11m | appears to be some icing | "
WXT_30m | not operational | "
Sonic_2m | some icing built up on the arms | "
Sonic_6m | some icing built up on the arms | "
Sonic_11m | some icing built up on the arms | "
Sonic_30m | not operational | "
Licor | some remnant icing on the arms, however the windows were clear and the red light was solid. Cleaned with a warm finger anyway. | "
SR50 | mostly ice free | "
IRT | cleared a little ice from the side of the body, but should not impact measurements. | "
Flux Plates | NR | "
CLASP | cleared some ice from the inlet. | "
SPCs (multiple) | NR | "
Data Files
Clasp02m |  | hourly files with size 3,095kB
CR1000X |  | daily file with size ~22500kB,
Licor02m |  | hourly files, raw with size ~3700-4000kB, stats with size ~19kB
Metek02m |  | hourly files, raw with size ~5280kB, stats with size 25kb
Metek06m |  | hourly files, raw with size ~5280kB, stats with size 25kb
Metek10m |  | hourly files, raw with size ~5280kB, stats with size 25kb
Metek30m |  | hourly files, raw with size ~5280kB, stats with size 25kb
Ozone |  | ????
Picarro |  | hourly files with size 3586kB

## ASFS40 (L1)
Atmospheric Surface Flux Stations in the Distributed Network
Note: Use this form for daily, remote assessments of ASFS. Use dedicated ASFS Site Visit form for site visits
Explanation
Communication System Status
Comms System | radio | [Radio, Iridium, or None] Indicate which forms of communication have been active
Quality | good | Provide any information on the quality of connection, how this impacts data flow
Power System Status
Operational Status | good | [Good, Questionable, Unknown]
Assessment of Exhaust | NR | Any notes about the buildup of ice that may be visible from the camera
Fuel Available | NR | Current remaining fuel available
Comments | Not much good data making it through for the EFOY. The measurements that do make it through show reasonable 24V. 12V is good. | Any other assessment of the power system, including quality of battery voltage/charge, operation of EFOY, etc.
Conditions
Structural | NR | Any useful information on the structural integrity of the ASFS
Position/Orientation/Level | no significant change | Any useful information on position; i.e., rapid changes in GPS position, GPS orientation, or sonic-inclinometer level
Surface | NR | Any useful information on local surface conditions from camera
Sky-Weather | NR | Any useful information on weather or sky conditions
Instrument/Measurement Assessment
LWD | mostly cloudy, with LWD>LWU for a couple hours early in the day. Also a very periods of partial clearing. Fan good. | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment from imagery if possible.
LWU | consistent with LWD and Tsurf. Fan good. | "
SWD | fan an dheat good | "
SWU | fan and heat good | "
IRT | Tsurf warmer than T_2m for most of the time by up to 0.8C, except for when the clouds thin at which point Tsurf gets colder than T_2m by up to 1.2C. Tsurface cools and warms before T_2m. | "
SR50 | Data appears to have changed into a more discretized mode with bigger steps between measurements | "
PTU | big mump in RH at about 14:00 UTC, could be exhaust although it does correspond with a slight perturbation in the pressure | "
Sonic | light and variable winds. | "
Licor | NA | "
Flux Plates | showing similar varability but offset by 4 W/m2 | "
Camera_Rad | NR | Any pictures available?
Camera_EFOY | NR | Any pictures available?
Data Logger | mainscan<4000. Period o lower values in the late middle part of the day. Ptemp=2-8C | Main scan time. Reasonable temperature range? Fans working?
Data Files
ASFS_40_slow | good | daily files of ~1120kB
ASFS_40_met | good | daily files of ~1740kB
ASFS_40_fast | good | daily files of >100MB
ASFS_40_Iridium_sumfile | no iridium | accumulating file with time stamp [include time stamp]

## ASFS30 (L2)
Atmospheric Surface Flux Stations in the Distributed Network
Note: Use this form for daily, remote assessments of ASFS. Use dedicated ASFS Site Visit form for site visits
Explanation
Communication System Status
Comms System | none | [Radio, Iridium, or None] Indicate which forms of communication have been active
Quality | NR | Provide any information on the quality of connection, how this impacts data flow
Power System Status
Operational Status | Down due to bear attack | [Good, Questionable, Unknown]
Assessment of Exhaust | NR | Any notes about the buildup of ice that may be visible from the camera
Fuel Available | NR | Current remaining fuel available
Comments | NR | Any other assessment of the power system, including quality of battery voltage/charge, operation of EFOY, etc.
Conditions
Structural | NR | Any useful information on the structural integrity of the ASFS
Position/Orientation/Level | NR | Any useful information on position; i.e., rapid changes in GPS position, GPS orientation, or sonic-inclinometer level
Surface | NR | Any useful information on local surface conditions from camera
Sky-Weather | NR | Any useful information on weather or sky conditions
Instrument/Measurement Assessment
LWD | NR | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment from imagery if possible.
LWU | NR | "
SWD | NR | "
SWU | NR | "
IRT | NR | "
SR50 | NR | "
PTU | NR | "
Sonic | NR | "
Licor | NR | "
Flux Plates | NR | "
Camera_Rad | NR | Any pictures available?
Camera_EFOY | NR | Any pictures available?
Data Logger | NR | Main scan time. Reasonable temperature range? Fans working?
Data Files
ASFS_30_slow | system down | daily files of ~1120kB
ASFS_30_met | system down | daily files of ~1740kB
ASFS_30_fast | system down | daily files of >100MB
ASFS_30_Iridium_sumfile | system down | accumulating file with time stamp [include time stamp]

## ASFS50 (L3)
Atmospheric Surface Flux Stations in the Distributed Network
Note: Use this form for daily, remote assessments of ASFS. Use dedicated ASFS Site Visit form for site visits
Explanation
Communication System Status
Comms System | iridium | [Radio, Iridium, or None] Indicate which forms of communication have been active
Quality | good | Provide any information on the quality of connection, how this impacts data flow
Power System Status
Operational Status | good | [Good, Questionable, Unknown]
Assessment of Exhaust | NR | Any notes about the buildup of ice that may be visible from the camera
Fuel Available | NR | Current remaining fuel available
Comments | EFOY profile is consistent. On for most of the day. 24V is good. | Any other assessment of the power system, including quality of battery voltage/charge, operation of EFOY, etc.
Conditions
Structural | NR | Any useful information on the structural integrity of the ASFS
Position/Orientation/Level | no change in inclination, but the system continues to rotate towards 239 from 245 a couple days ago. | Any useful information on position; i.e., rapid changes in GPS position, GPS orientation, or sonic-inclinometer level
Surface | NR | Any useful information on local surface conditions from camera
Sky-Weather | NR | Any useful information on weather or sky conditions
Instrument/Measurement Assessment
LWD | mostly cloudy. LWD>LWU early. Some slightly clearing periods. Fan good. | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment from imagery if possible.
LWU | consistent with LWD and Tsurf. Fan good. | "
SWD | fan and heat good | "
SWU | fan and heat good | "
IRT | Looks reasonable | "
SR50 | some erratic measurements towards the end of the day. | "
PTU | Looks reasonable. | "
Sonic | consistent data with mild winds. | "
Licor | NA | "
Flux Plates | Larger values have diminished as the air temperatures warmed somewhat over the day. | "
Camera_Rad | NR | Any pictures available?
Camera_EFOY | NR | Any pictures available?
Data Logger | mainscan=120, Ptemp=-5 to 2C | Main scan time. Reasonable temperature range? Fans working?
Data Files
ASFS_50_slow | no radio | daily files of ~1120kB
ASFS_50_met | no radio | daily files of ~1740kB
ASFS_50_fast | no radio | daily files of >100MB
ASFS_50_Iridium_sumfile | good | accumulating file with time stamp [include time stamp]

## ARM
ARM Facility - General Data Oversight
Explanation
Measurement Assessments |  | Assessment of data for scientific quality. Do the measurements look reasonable compared to conditions? Are the measurements capturing the key features?
Clouds
KAZR | good
MWACR
Ka-SACR
X-SACR
HSRL
MPL | good, with data gap 07:00-08:30 UTC
Ceilometer | good
MWR2C | good
MWR3C
TSI
Precipitation
LDIS/PARS2 (ice)
LDIS/PARS2 (ship)
PWD (ice)
PWD (ship)
Siphon rain gauge (ship)
ORG (ice)
WBRG/Pluvio (ice)
Winds
BSRWP
DL
Radiation
AERI
ShipRad - PIR, PSP, SPN1
SkyRad - PSP, PIR, SPN1 (ice)
GndRad - PIR, PSP (ice)
MFR (ice)
CSPHOT
FRSR
Aerosol
AOSMet
CO / Purge
O3
CPC
CPCU
SMPS
HTDMA
UHSAS
CCNC
ACSM
SP2
PSAP
Nephelometer
INP filters
Operated ARMs on ice instrumentation on generator starting at about 4 UTC.
Data Requests |  | Any facilitation of MOSAiC data users

## Guests
Guest and Partner Instruments
Explanation
Power Distribution Unit - Met Hut (PFS)
Operational status | 11:00 UTC started a generator and connected one of the UPS. It responded and started charging batteries.
Power balance on two PDUs? | Only one operational
Back-up generator test (/month) | Backup generator #1 is operational. The second has not been operated yet.
Precip VISSS - Met City (Maahn)
not operational
Doppler Lidar - Met City (Brooks)
not operational
Doppler Lidar - Polarstern (Brooks)
Brooks onsite
Doppler Lidar - Polarstern (Heinemann)
Brooks onsite
Sodar - Met City (Brooks)
not operational
Radiation Suite - Met City (Pirazzini)
not operational