## General
Surface Flux Team Daily Checks
If nothing to report use NR
Explanation
General
Date: | 20191210 | [YYYYMMDD]
Operator1: | Matt Shupe, Dave Costa | [name]
Operator2: | Ola Persson | [name]
Status of Operators: | NR | Any concerns or limitations regarding operators? Health, injury, etc.
General Weather Conditions: | mild, with some wind | Brief weather synopsis, especially weather that would impact operations
General Sea Ice Conditions: | No major changes | Brief local sea ice synopsis, especially conditions that would impact operations
Special Operations: | Met City visit: 06:15-08:00 | Any daily activities or special operations that would possibly impact daily schedule or operations
Camp Orientation: | No change | Any information about the general composition, orientation, and/or layout of camp that would impact measurements
Command Center
Utilities: | Good | [comments or NR] Status on power connection, power panel, heaters, lights, network connection
Physical Integrity: | Good | [comments or NR] Any damage, ingress of water/snow, snow/water management, other structural concerns
DAQ Computer: | Good
Workstation Computer: | Good
Communications Mode(s): | radio, iridium. Cable to Met City still down. Wireless LAN to Met City up. | [Radio, Iridium, Cable, and/or None] Which comms systems are being used?
Communications Quality: | Good for the operational systems | Any information about operation of the radio and/or Iridium modems. What is the quality of connection? Any actions that were taken to modify the quality
Remote Mounts
Iridium Antenna - ARM Container | good | Physical system assessment. Describe any actions taken such as cleaning or adjustments
Radio modem/antenna(s) - Crow's Nest | NR | Physical system assessment. Describe any actions taken such as cleaning or adjustments
Time-Lapse Camera - Crow's Nest | NA | Operational? Comment on visual assessment, and any actions taken. Data downloaded?
ACTION ITEMS
List things that need to be accomplished in the coming days, developing issues, etc.
Supplies
Ongoing list of supplies needed in the field; carry this forward until supplies have been obtained

## Met Tower
Met Towers at Met City (11m tower and 30m mast)
Explanation
Local Conditions
Ice and Snow Conditions | No change from yesterday | Ice thickness, cracks, snow drifting, ablation or any other relevant conditions for operations.
Ice and Snow Management | Managed some snow around the Met Hut to mitigate future drifting | Anything required for physical maintenance of local area
Utility Status
Power line connectivity | up | [up or down, comment on stability]
Network connectivity | fiber down, wireless ship LAN up | [up or down, comment on stability]
Utility cables | good. Did some management of the Stanton, FMI poles | [good or in need of maintenance]. Utility cables on poles or maintained routinely on surface, connections secure, strain relief intact, etc.
Control box | good. Cut the power for a short time to test power system | Any comments on the operations of the computer, data logger, power distribution, or other components. Temperature range and stability inside.
SPC Control box | good. Everything appears to be functional | Any comments on Frey's power and control box.
Structural Assessment - 10m
Tower Base | Stable | Stability of the base, including any potential ablation
Guy Anchors | Stable | Any sign of anchor instability?
Guy Tightness | good. Tightness is slightly uneven due to winds | Even tightness? If not adjust.
Structural Integrity | solid | Any issues identified with structural components? Bolts, clamps, cables checked for tightness?
Plumbness | NR | Assess any shifts in plumbness of tower, possibly using manual measurements and/or sonic inclinometers
Boom Level / Stability | NR | Assess stability and level of booms
Cabling | NR | Secure with no rubbing.
Structural Assessment - 30m
Tower Base | stable | Stability of the base, including any potential ablation
Guy Anchors | stable | Any sign of anchor instability?
Guy Tightness | NR | Even tightness? If not adjust.
Structural Integrity | appears to be good | Any issues identified with structural components? Bolts, clamps, cables checked for tightness?
Plumbness | appears to be good | Assess any shifts in plumbness of tower, possibly using manual measurements and/or sonic inclinometers
Cabling | NR | Secure with no rubbing.
Instrument/Measurement Assessment
Daily Cleaning | Met City visit 06:15-08:00. Tower cleaning not needed on most instruments except for Licor. Short interruption of power on the tower instruments. Shut down the tower computer cleanly, but then it needed the button to turn it back on. On first attempt the "30m" sonic didn't come back, but upon restart of the software everything came back. | Include the timing and any details related to cleaning
PTU_2m | ice free | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment, and any actions taken.
HMT_6m | ice free | "
HMT_10m | ice free | "
WXT_30m | operational but difficult to assess icing. | "
Sonic_2m | ice free | "
Sonic_6m | ice free | "
Sonic_10m | ice free | "
Sonic_30m | operational but difficult to assess icing. | "
Licor | small amount of ice buildup. Cleaned with ethanol. Then wiped dry with the soft part of glove with finger. | "
SR50 | ice free. Short lateral drift continues to evolve just between the viewed surface and the tower; this might be impacting the snow depth measurements. | "
IRT | ice free | "
Flux Plates | NR | "
CLASP | ice free | "
SPCs (multiple) | ice free | "
Data Files
Clasp02m |  | hourly files with size 3,095kB
CR1000X |  | daily file with size ~22500kB,
Licor02m |  | hourly files, raw with size ~3700-4000kB, stats with size ~19kB
Metek02m |  | hourly files, raw with size ~5280kB, stats with size 25kb
Metek06m |  | hourly files, raw with size ~5280kB, stats with size 25kb
Metek10m |  | hourly files, raw with size ~5280kB, stats with size 25kb
Metek30m |  | hourly files, raw with size ~5280kB, stats with size 25kb
Ozone |  | ????
Picarro |  | hourly files with size 3586kB

## ASFS40 (L1)
Atmospheric Surface Flux Stations in the Distributed Network
Note: Use this form for daily, remote assessments of ASFS. Use dedicated ASFS Site Visit form for site visits
Explanation
Communication System Status
Comms System | radio | [Radio, Iridium, or None] Indicate which forms of communication have been active
Quality | ok but getting better | Provide any information on the quality of connection, how this impacts data flow
Power System Status
Operational Status | good | [Good, Questionable, Unknown]
Assessment of Exhaust | NR | Any notes about the buildup of ice that may be visible from the camera
Fuel Available | NR | Current remaining fuel available
Comments | Little EFOY data but some voltages coming in at 28+V. Ptemp curve continues to look a bit strange, almost like the ventilation is behaving strangely. But overall it appears that the EFOy system is working well. | Any other assessment of the power system, including quality of battery voltage/charge, operation of EFOY, etc.
Conditions
Structural | NR | Any useful information on the structural integrity of the ASFS
Position/Orientation/Level | no significant change | Any useful information on position; i.e., rapid changes in GPS position, GPS orientation, or sonic-inclinometer level
Surface | NR | Any useful information on local surface conditions from camera
Sky-Weather | NR | Any useful information on weather or sky conditions
Instrument/Measurement Assessment
LWD | mostly clear with a late morning cloudy period. Fan good | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment from imagery if possible.
LWU | increase during the mid-morning cloud-induced warming. Fan good. | "
SWD | fan and heat good | "
SWU | fan and heat good | "
IRT | 2+C inversion below 2m. | "
SR50 | NR | "
PTU | downward spike in RH at 22:00 UTC. Otherwise data looks reasonable | "
Sonic | reasonable measurements with variable winds over the day. | "
Licor | NA | "
Flux Plates | two plates tracking but offset by 9 W/m2 | "
Camera_Rad | NR | Any pictures available?
Camera_EFOY | NR | Any pictures available?
Data Logger | mainscan=3600-4000, Ptemp=-7 to 6C | Main scan time. Reasonable temperature range? Fans working?
Data Files
ASFS_40_slow | good | daily files of ~1120kB
ASFS_40_met | good | daily files of ~1740kB
ASFS_40_fast | turned off to preserve radio | daily files of >100MB
ASFS_40_Iridium_sumfile | no iridium | accumulating file with time stamp [include time stamp]
Extra Comments on Site Visit

## ASFS30 (L2)
Atmospheric Surface Flux Stations in the Distributed Network
Note: Use this form for daily, remote assessments of ASFS. Use dedicated ASFS Site Visit form for site visits
Explanation
Communication System Status
Comms System | radio | [Radio, Iridium, or None] Indicate which forms of communication have been active
Quality | marginal but appears to be getting better | Provide any information on the quality of connection, how this impacts data flow
Power System Status
Operational Status | Good | [Good, Questionable, Unknown]
Assessment of Exhaust | NR | Any notes about the buildup of ice that may be visible from the camera
Fuel Available | NR | Current remaining fuel available
Comments | EFOY profile similar to what it has been. Peak voltage right at the end of the day reached up to 26.9V. | Any other assessment of the power system, including quality of battery voltage/charge, operation of EFOY, etc.
Conditions
Structural | NR | Any useful information on the structural integrity of the ASFS
Position/Orientation/Level | no significant change | Any useful information on position; i.e., rapid changes in GPS position, GPS orientation, or sonic-inclinometer level
Surface | NR | Any useful information on local surface conditions from camera
Sky-Weather | NR | Any useful information on weather or sky conditions
Instrument/Measurement Assessment
LWD | clear sky for the whole day except a period in the morning of cloudiness. Fan good. | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment from imagery if possible.
LWU | temporary increase during the cloudy period, but otherwise lower. Fan good. | "
SWD | fan and heat good | "
SWU | fan and heat good. Spikey nature of fantach continues but in the second half of the day it went away. | "
IRT | a couple of spikes in the data that will need to be filtered out. Otherwise, near surface inversion of 2.5C for much of the day. | "
SR50 | a spike in the data will need to be filtered out. | "
PTU | reasonable data | "
Sonic | reasonable data | "
Licor | NA | "
Flux Plates | follow each other closely | "
Camera_Rad | NR | Any pictures available?
Camera_EFOY | NR | Any pictures available?
Data Logger | mainscan=~2700-3300, Ptemp=-6 to 6C | Main scan time. Reasonable temperature range? Fans working?
Data Files
ASFS_30_slow | good | daily files of ~1120kB
ASFS_30_met | good | daily files of ~1740kB
ASFS_30_fast | turned off to preserve radio | daily files of >100MB
ASFS_30_Iridium_sumfile | no iridium | accumulating file with time stamp [include time stamp]
Extra Comments on Site Visit

## ASFS50 (L3)
Atmospheric Surface Flux Stations in the Distributed Network
Note: Use this form for daily, remote assessments of ASFS. Use dedicated ASFS Site Visit form for site visits
Explanation
Communication System Status
Comms System | iridium | [Radio, Iridium, or None] Indicate which forms of communication have been active
Quality | good | Provide any information on the quality of connection, how this impacts data flow
Power System Status
Operational Status | good | [Good, Questionable, Unknown]
Assessment of Exhaust | NR | Any notes about the buildup of ice that may be visible from the camera
Fuel Available | NR | Current remaining fuel available
Comments | little EFOY data. See some 28+V coming in with charge current up to 4 Amps. Suggestion is that the system is operating well. | Any other assessment of the power system, including quality of battery voltage/charge, operation of EFOY, etc.
Conditions
Structural | NR | Any useful information on the structural integrity of the ASFS
Position/Orientation/Level | no significant change | Any useful information on position; i.e., rapid changes in GPS position, GPS orientation, or sonic-inclinometer level
Surface | NR | Any useful information on local surface conditions from camera
Sky-Weather | NR | Any useful information on weather or sky conditions
Instrument/Measurement Assessment
LWD | mid-morning cloudiness in otherwise clear skies. Fan good | Comment on data quality (in range, reasonable variability, expected correspondance with other measurements, etc.), visual assessment from imagery if possible.
LWU | increase in Lwu when surface warms with clouds. Fan good. | "
SWD | fan and heat good | "
SWU | fan and heat good | "
IRT | near surface inversion of 2C. | "
SR50 | NR | "
PTU | reasonable measurements | "
Sonic | pretty good data recovery on iridium, showing variable winds over the day. | "
Licor | NA | "
Flux Plates | two plates tracking each other with large offset of up to 30W/m2 | "
Camera_Rad | NR | Any pictures available?
Camera_EFOY | NR | Any pictures available?
Data Logger | mainscan=120, Ptemp=-9 to 1C | Main scan time. Reasonable temperature range? Fans working?
Data Files
ASFS_50_slow | no radio | daily files of ~1120kB
ASFS_50_met | no radio | daily files of ~1740kB
ASFS_50_fast | no radio | daily files of >100MB
ASFS_50_Iridium_sumfile | good | accumulating file with time stamp [include time stamp]
Extra Comments on Site Visit

## ARM
ARM Facility - General Data Oversight
Explanation
Measurement Assessments | Assessment of data for scientific quality. Do the measurements look reasonable compared to conditions? Are the measurements capturing the key features?
Clouds
KAZR
MWACR
Ka-SACR
X-SACR
HSRL
MPL
Ceilometer
MWR2C
MWR3C
TSI
Precipitation
LDIS/PARS2 (ice)
LDIS/PARS2 (ship)
PWD (ice)
PWD (ship)
Siphon rain gauge (ship)
ORG (ice)
WBRG/Pluvio (ice)
Winds
BSRWP
DL
Radiation
AERI
ShipRad - PIR, PSP, SPN1
SkyRad - PSP, PIR, SPN1 (ice)
GndRad - PIR, PSP (ice)
MFR (ice)
CSPHOT
FRSR
Aerosol
AOSMet
CO / Purge
O3
CPC
CPCU
SMPS
HTDMA
UHSAS
CCNC
ACSM
SP2
PSAP
Nephelometer
INP filters
Data Requests | Any facilitation of MOSAiC data users

## Guests
Guest and Partner Instruments
Explanation
Power Distribution Unit - Met Hut (PFS)
Operational status | Down. Tried to start up both of the UPS units but they would not start.
Power balance on two PDUs?
Back-up generator test (/month)
Precip VISSS - Met City (Maahn)
All checks good. Downloaded the data and will upload o MCS
Doppler Lidar - Met City (Brooks)
Operational.
Doppler Lidar - Polarstern (Brooks)
Operational. Brooks onsite.
Doppler Lidar - Polarstern (Heinemann)
Operational. Brooks onsite.
Sodar - Met City (Brooks)
Operational. Short power interruption but the system came back up just fine.
Radiation Suite - Met City (Pirazzini)
Domes rimed up. Cleaned during Met City visit. Down welling Long wave was iced over. Cleaned with ethanol.