## Sheet1
Sdcard | Photos |  |  |  |  | Leg 1
|  | Log of remote site visits (refer to daily logs for more details of site visit). |  |  |  | Leg 2
Leg 2 - other team visits
|  | Date | Site | Brief summary of activities | Fuel burn rate | Leg 3
|  | 20191005 | L1 (ASFS40) | Set up site . Extensive documentation, incl. snow depth.
|  | 20191007 | L2 (ASFS30) | Set up site. Extensive documentation, incl. snow depth.
|  | 20191010 | L3 (ASFS50) | Set up site Oct 9-10. Extensive documentation, incl. snow depth.
|  | 20191015 | L1 (ASFS40) | Re-visit of L1 from Federov, remove Licor from operation | 1.73 L/day with Licor
|  | 20191016 | L2 (ASFS30) | Re-visit of L2 from Federov, remove Licor from operation | 1.43 L/day with Licor
|  | 20191106 | L2 (ASFS30) | Change LWD ventilator. Fuel change. Basic maintenance, incl. snow depth measurements | 1.00 L/day w/o Licor
|  | 20191110 | L1 (ASFS40) | Swap dead EFOY. Fuel change. Basic maintenance. | 1.19 L/day w/o Licor
|  | 20191114 | L3 (ASFS50) | Fuel change. Basic maintenance, incl snow depth measurements. | 1.63 L/day with Licor
|  | 20191114 | L2 (ASFS30) | Assess failure. Found bear damage to the station. Retrieved the EFOY and batteries.
|  | 20191122 | L2 (ASFS30) | Remove damaged equipment from bear damage. EFOY exhaust system, IRT, SWD and a part of the RF antenna cable. Installed replacement sonic anemometer
|  | 20191130 | L2 (ASFS30) | Re-install full EFOY system including rebuilt exhaust, IRT cable, SWD cable, radio cable, swapped SWU instruments, straightened Met mount, basic maintenance, No fuel swap needed.
|  | 20191206 | L2 (ASFS30) | Fix the IRT (found blown fuse), measure snow depth, tape sonic box connectors.
|  | 20191206 | L3 (ASFS50) | Fuel change. Basic maintenance. Assess and change PTU system. Assess SR50 (no issue observed). Turn off power to licor. | 1.78 L/day with Licor powered
|  | 20191207 | L1 (ASFS40) | Fuel change. Basic maintenance. Taped the sonic box connectors. Some exhaust beard. Snow in SR50. IRT bracket was loose-tightened. Snow depth measurements. No other issues. | 1.37 L/day w/o Licor
| x | 20191222 | L1 (ASFS40) | First flight to L-site for Gallagher/Uttal and helicopter crew. Landing lights were replaced. There is a new ruggedized battery system being implemented that should facitate lifetime. ASFS40 was visually checked. Soncs ad radiometers are clear. Large exhaust ice sculpture removed - does not seem to be impairing exhaust
|  | 20191223 | L2 (ASFS30) | Replace Landing Lights
|  | 20191223 | L3 (ASFS50) | Replace Landing Lights
|  | 20191223 | L3 (ASFS50) | Second flight of the day to L3 - Visual system check (Gallagher). Cleared exhaust system ice build up.
|  | 20191223 | L2 (ASFS30) | Second and third flight of the day to L2 - Fuel replacement - Gallagher and logistics bear guard, Haas as methanol minder. Visual inspection and cleared exhaust system ice build up
| x | 20191228 | L1 (ASFS40) | Replaced SD card. Cleared exhaust ice
|  | 20191228 | L2 (ASFS30) | Replaced SD card. Cleared exhaust ice
|  | 20191228 | L3 (ASFS50) | Replaced SD card. Cleared exhaust ice
| x | 20200106 | L1 (ASFS40) | Refueled. Site Survey by B.Rabe and M. Gallagher.
|  | 20200106 | L3 (ASFS50) | Refueled. Replaced Sdcard. Site Survey by B. Rabe and M. Gallagher.
|  | 20200106 | L2 (ASFS30) | Site Survey by B. Rabe and M. Gallagher.
|  | 20200108 | L3 (ASFS50) | Site Visit to disconnect radio. No site visit folders created since no photos or Sdcard data.
| m | 20200111 | L1 (ASFS40) | Reset Fuel switch
|  | 20200111 | L3 (ASFS50) | Upload logger updates (comms issue). Retreived SD card. Snow measurement 15 cm
| x | 20200117 | L1 (ASFS40) | Team Ice visited and cleaned exhaust and took photos
|  | 20200117 | L3 (ASFS50) | Team Ice visited and cleaned exhaust and took photos - also noted that there is a narrow crack which is less than a meter about 4-5 meters from L3. The crack is already frozen over and filling with snow.
|  | 20200120 | L2 (ASFS30) | Refuel and data card swap. Standard Maintenance and Checks.
| x | 20200120 | L1 (ASFS40) | Data card swap. Standard Maintenance Checks.
|  | 20200120 | L3 (ASFS50) | Data card swap. Standard Maintenance Checks.
|  | 20200122 | L2 (ASFS30) | Team ICE journeyed to L2 by Skidoo. Cleaned Exhaust
|  | 20200125 | L3 (ASFS50) | Welness Check on L3 that had no radio pings and showed a last inside box temperature of -19 degree C. The system was not immediately in danger from near by cracks/leads but was down. Apparently due to complete exhaust freezeup/blockage. Sdcard, Batteries and exhaust system retrieved.
| m | 20200130 | L3 (ASFS50) | EFOY SWAP for L3. The L3 Batteries were recovered on 1/25/ but there was not a successful warming and recharging. The warming location was the helicopter hanger and both the battaries ruptured while warming. Consequently the spare batteries had to be used: This has creating an issue with connections as the spares (ahem) do not have the same size posts. How it went: MIchael and MIsha (bear guard) flew out and set up camp. The helicopter came back and got Taneil and the toasty EFOY set up. Michael made the swap. We hung around for a couple of hours to see if everything came on and they EFOY kept running. Everything looked good before we left except the sonic. It may have still been warming up. NOTE: Chris got some spare-spare batteries before leaving Tromso
| m | 20200202 | L1 (ASFS40) | Sdcard retrieval. Blocked section of intake vent - IR camera showed it to be a large heat loss source. Box temperatures have been out of range cold.
| m | 20200206 | L3 (ASFS50) | ASFS-50 found tipped up in ridge. Sonic trashed. Boom bent and vaisala intrument looked damaged. Unable to open box to retreive SD card. Unable to fully assess as helicopter called back to ship
| m | 20200209 | L3 (ASFS50) | System Secured with anchors to ice. Large Lead has now opened up next to the ridge that ASFS-50 is buried in. SD card retreived. Batteries retreived.
| m | 20200210 | L3 (ASFS50) | System back on skids. Moved to flat area near surviving buoys - there have also been 2 buoy casualties at L3. Licor and Vaisala instruments recovered. EFOY recovered. Radiometers, SR50, GPS, Radio and IRT look intact. Radio antenna and Interface box appears intact. Frame appears intact. Slid nicely to it's new location.
|  | 20200215 | L1 (ASFS40) | Sdcard retrieval. Refueled. Likely automatically switched tanks.
|  | 20200215 | L2 (ASFS30) | Sdcard retrieval. Added foam to intake vent. Refueled.
|  | 20200222 | L2 (ASFS30) | Refuel. Note: left the lights on. :-(
|  |  | L2 (ASFS30) | Cox, Gallagher, Brossier snow machines. EFOY Swap. Sdcard retrieved. Charging when we left…death in 36 hours. -39 to -42