## June 27
MOSAiC Data Acquisition Timestamp Benchmark Test June 27, 2019 00:00 UTC - June 27, 2019 14:00 UTC (datalogger ingest active; remote-desktop connection active; clock synchronization active; remote file transfers @ )
Instrument | Communication Protocol | Expected # of Samples | Ingested # of Samples | Sample Difference | Lagged Samples | Sampling interval offsets (5-millisecond bins) listed as a percentage of the total number of ingested samples (% >= 1 highlighted in bold red)
[6 empty] | -47.5 | 42.5 | -37.5 | -32.5 | -27.5 | -22.5 | -17.5 | -12.5 | -7.5 | -2.5 | 0 ms | 2.5 | 7.5 | 12.5 | 17.5 | 22.5 | 27.5 | 32.5 | 37.5 | 42.5 | 47.5
CLASP (02 m) | RS-422 serial | 540000 | 540077 | 77 | 0 | [8 empty] | 60.09 | 0.006 |  |  | 24.97 | 14.94
Licor (02 m) | TCP socket | 1008000 | 1079980 | -20 | 0 | [9 empty] | 80 | 0.0002 |  |  | 19.99 |  |  | 0.0059 |  |  | 0.003
Metek (02 m) | TCP socket | 1008000 | 1080005 | 5 | 0 | [9 empty] | 80 | 0.0002 |  |  | 20
Metek (06 m) | TCP socket | 1008000 | 1080006 | 6 | 0 | [9 empty] | 80 | 0.0001 |  |  | 20
Metek (10 m) | TCP socket | 1008000 | 1080009 | 9 | 0 | [8 empty] | 0.0001 | 80 | 0.0001 | 0.0001 |  | 20
Metek (30 m) | RS-422 serial | 1008000 | 1080010 | 10 | 0 | [7 empty] | 0.0009 | 0.0006 | 80 | 0.001 | 0.0023 | 0.0027 | 19.99
Ozone | RS-485 serial | 504000 | 540010 | 10 | 0 | [5 empty] | 10.09 | 0.0002 |  | 45.85 | 0.0004 |  |  | 23.7 | 14.34 |  | 0.4752 | 5.55
Picarro | TCP socket | ? | 849732 | ? | 0 | [6 empty] | 2 | 11.48 |  |  |  | 80.41 | 0.0001 |  | 3.37 |  |  | 0.6649 |  |  | 0.3631 #
Notes: @ data and log file transfers from the ship computer occurred every minute on an ethernet channel that was isolated from the instrument data streams
? raw sample rate is approximately 15.8 Hz for the current string length
# remaing Picarro samples had differences between +50-500 milliseconds

## June 26
MOSAiC Data Acquisition Timestamp Benchmark Test June 26, 2019 00:00 UTC - June 26, 2019 14:00 UTC (datalogger ingest active; remote-desktop connection active; clock synchronization active; hourly remote file transfers @ )
Instrument | Communication Protocol | Expected # of Samples | Ingested # of Samples | Sample Difference | Lagged Samples | Sampling interval offsets (5-millisecond bins) listed as a percentage of the total number of ingested samples (% >= 1 highlighted in bold red)
[6 empty] | -47.5 | 42.5 | -37.5 | -32.5 | -27.5 | -22.5 | -17.5 | -12.5 | -7.5 | -2.5 | 0 ms | 2.5 | 7.5 | 12.5 | 17.5 | 22.5 | 27.5 | 32.5 | 37.5 | 42.5 | 47.5
CLASP (02 m) | RS-422 serial | 504000 | 504072 | 72 | 0 | [8 empty] | 60.09 | 0.006 |  |  | 24.98 | 14.92
Licor (02 m) | TCP socket | 1008000 | 1007982 | -18 | 0 | [5 empty] | 0.004 | 0.0743 |  |  | 79.85 | 0.0001 | 0.002 |  | 20.07 |  |  | 0.0027 |  |  | 0.003
Metek (02 m) | TCP socket | 1008000 | 1008002 | 2 | 0 | [9 empty] | 80 |  | 0.001 |  | 20 | 0.001
Metek (06 m) | TCP socket | 1008000 | 1008007 | 7 | 0 | [9 empty] | 80 | 0.002 |  |  | 20 |  |  | 0.0001
Metek (10 m) | TCP socket | 1008000 | 1008009 | 9 | 0 | [6 empty] | 0.0001 |  |  | 80 | 0.0001 |  |  | 20
Metek (30 m) | RS-422 serial | 1008000 | 1008009 | 9 | 0 | [7 empty] | 0.0049 | 0.0077 | 79.97 | 0.001 | 0.0175 | 0.0145 | 19.99
Ozone | RS-485 serial | 504000 | 504008 | 8 | 0 | [5 empty] | 10.04 |  |  | 45.93 | 0.0006 |  |  | 23.85 | 14.16 |  | 0.5222 | 5.49
Picarro | TCP socket | ? | 796535 | ? | 0 | [6 empty] | 2 | 11.49 | 0.0001 | 0.0003 |  | 80.55 | 0.0003 |  | 3.36 |  |  | 0.6395 |  |  | 0.3640 #
Notes: @ file transfers by the ship computer occurred at 2 minutes past the top of the hour on the same ethernet channel as the instrument data acquisition
+ between the range +55 and +65 milliseconds
? raw sample rate is approximately 15.8 Hz for the current string length
# remaing Picarro samples had differences between +50-410 milliseconds

## June 14
MOSAiC Data Acquisition Timestamp Benchmark Test June 14, 2019 00:00 UTC - June 14, 2019 13:00 UTC (datalogger ingest active; remote-desktop connection active; clock synchronization active; hourly remote file transfers @ )
Instrument | Communication Protocol | Expected # of Samples | Ingested # of Samples | Sample Difference | Sampling interval offsets (5-millisecond bins) listed as a percentage of the total number of ingested samples (% >= 1 highlighted in bold red)
[5 empty] | -47.5 | 42.5 | -37.5 | -32.5 | -27.5 | -22.5 | -17.5 | -12.5 | -7.5 | -2.5 | 0 ms | 2.5 | 7.5 | 12.5 | 17.5 | 22.5 | 27.5 | 32.5 | 37.5 | 42.5 | 47.5
CLASP (02 m) | RS-422 serial | 468000 | 468065 | 65 | [5 empty] | 0.0017 |  | 0.0006 | 59.93 | 0.0013 |  | 0.0009 | 26.07 | 13.99 |  |  | 0.0006
Licor (02 m) | TCP socket | 936000 | 935946 | -54 # | [5 empty] | 0.001 | 0.1143 |  | 0.0018 | 79.68 | 0.0003 | 0.004 | 0.004 | 20.19 |  |  | 0.0046
Metek (02 m) | TCP socket | 936000 | 935996 | -4 | [8 empty] | 0.0032 | 79.92 | 0.0001 | 0.0011 | 0.0003 | 20.08
Metek (06 m) | TCP socket | 936000 | 935995 | -5 | [6 empty] | 0.0009 |  | 0.003 | 79.92 | 0.001 | 0.0015 | 0.0005 | 20.08 |  |  | 0.0001
Metek (10 m) | TCP socket | 936000 | 935999 | -1 | [6 empty] | 0.0009 |  | 0.0032 | 79.92 | 0.0005 | 0.0016 | 0.0005 | 20.08
Metek (30 m) | RS-422 serial | 936000 | 935999 | -1 | [6 empty] | 0.0001 | 0.0021 | 0.0059 | 79.91 |  | 0.0082 | 0.0065 | 20.07
Ozone | RS-485 serial | 468000 | 468005 | 5 | $ |  | 0.0141 |  |  | 0.0139 |  | 0.0013 | 59.81 | 0.0013 |  | 0.0011 | 26.09 | 14.04 |  | 0.0009 | 0.01 |  |  | 0.0145 | $
Picarro | TCP socket | ? | 765730 | ? | [6 empty] | 1.94 | 11.27 | 0.0003 | 0.0021 | 0.0034 | 84.58 | 0.001 |  | 1.64 |  |  | 0.1783 |  |  | 0.0804 #
Notes: @ file transfers by the ship computer occurred at 2 minutes past the top of the hour on the same ethernet channel as the instrument data acquisition
# Licor sample ingest/time-stamping lagged behind by one sample 1-5 times during each hourly file-transfer session; the lagged samples were not lost, but were timestamped outside of the 50 ms tolerence; the -54 difference represents a slower than advertised data rate
$ 46 Ozone samples had differences between +/- 50-90 milliseconds
? raw sample rate is approximately 16 Hz for the current string length
# 2347 Picarro samples had differences between +50-378 milliseconds

## June 9
MOSAiC Data Acquisition Timestamp Benchmark Test June 9, 2019 21:00 UTC - June 10, 2019 15:00 UTC (datalogger ingest active; remote-desktop connection inactive; clock synchronization inactive; hourly remote file transfers inactive; Picarro output stream incorrect)
Instrument | Communication Protocol | Expected # of Samples | Ingested # of Samples | Sample Difference | Sampling interval offsets (5-millisecond bins) listed as a percentage of the total number of ingested samples (% >= 1 highlighted in bold red)
[5 empty] | -47.5 | 42.5 | -37.5 | -32.5 | -27.5 | -22.5 | -17.5 | -12.5 | -7.5 | -2.5 | 0 ms | 2.5 | 7.5 | 12.5 | 17.5 | 22.5 | 27.5 | 32.5 | 37.5 | 42.5 | 47.5
CLASP (02 m) | RS-422 serial | 720000 | 720102 | 102 | [5 empty] | 0.0039 |  | 0.0007 | 59.93 | 0.0011 | 0.0001 | 0.0008 | 26.08 | 13.98 |  | 0.0001 | 0.0018
Licor (02 m) | TCP socket | 1440000 | 1439974 | -26 | [6 empty] | 0.0953 |  | 0.0015 | 79.74 | 0.0001 | 0.0008 | 0.0006 | 20.15 | 0.0001 |  | 0.0101 |  |  | 0.0001
Metek (02 m) | TCP socket | 1440000 | 1439998 | -2 | [8 empty] | 0.003 | 79.92 | 0.0005 | 0.0015 | 0.0005 | 20.08
Metek (06 m) | TCP socket | 1440000 | 1439995 | -5 | [6 empty] | 0.0039 | 0.0001 | 0.0026 | 79.91 | 0.0003 | 0.0012 | 0.0003 | 20.08 |  |  | 0.0003
Metek (10 m) | TCP socket | 1440000 | 1440002 | 2 | [8 empty] | 0.0035 | 79.92 | 0.0001 | 0.0012 | 0.0003 | 20.08
Metek (30 m) | RS-422 serial | 1440000 | 1440001 | 1 | [7 empty] | 0.0024 | 0.0064 | 79.9 | 0.0003 | 0.0088 | 0.0068 | 20.07
Ozone | RS-485 serial | 720000 | 720011 | 11 | 0.0032 * |  | 0.0026 |  |  | 0.0079 |  | 0.0006 | 59.84 | 0.0011 | 0.0001 | 0.0007 | 26.12 | 14.01 |  |  | 0.0044 |  |  | 0.0033 | 0.0022 *
Picarro | TCP socket | ? | 664574 | ? | [9 empty] | data stream currently in burst mode with latencies up to 1500 milliseconds; distribution N/A #
Notes: * 39 Ozone samples had differences between +/- 50-90 milliseconds
? data rate unknown; remote system maintenance also occurred at unknown times
# manufacturer working on system to configure/enable raw data stream