## Sheet1
Instruments for MOSAiC Flux measurements
Component | Model | Parameters | Installation | Frequency | Operation T Range | Power Requirements | Power Consumption, nominal | Power Consumption, peak | Power for heating | Comms | Data Rate | Comment | Cost | Application | More Comments
LW up/down | Hukseflux IR20-T2 | LWU, LWD | ASFS-2.5m | 1 sec | -40 to 80 C | 12 VDC | <0.1 W |  | 1.5 W optional |  |  |  | 4462 | ASFS | IR20-T2 with LM01 mounting plate and 5m cable.
| Hukseflux VU01 housing for IR20 |  |  |  |  | 12 VDC | 6.7 W (unheated) |  | 5 or 10 W options |  |  |  | 1345 | ASFS | VU01, with standard 5 m cable
| Kipp and Zonen SGR4 (digital) | LWU, LWD | ASFS-2.5m | 1sec | -40 to 80 C | 5-24 VDC | <0.1 W |  | none | RS485 Modbus |  | have analog option |  | ASFS
| Kipp and Zonen CVF4 ventilator |  |  |  |  | 12 VDC | 10.1 W (incl. 6 W heat) |  | 5 or 10 W options |  |  | Bulky! |  | ASFS
SW up/down | Hukseflux SR30-D1 (digital) | SWU, SWD | ASFS-2.5m | 1 sec | -40 to 80 C | 8-30 VDC | <2.3W (<0.1W in low power) |  | included | RS485 RTU / ModBus |  | can operate in low power mode with heater / ventilator off | 2695 | ASFS | SR30-D1-LM01-05. This includes the flat mounting bracket and standard 5-m cable length,
| Kipp and Zonen SMP22 (digital) | SWU, SWD | ASFS-2.5m | 1 sec | -40 to 80 C | 5-30 VDC | 0.1 W |  | none | RS485 Modbus |  | Also have analog systems |  | ASFS
| Kipp and Zonen CVF4 ventilator |  |  |  |  | 12 VDC | 10.8 W (w/ 5.5 W heater) |  | 5 or 10 W options |  |  | Bulky! |  | ASFS
SW diffuse/direct | Delta-T SPN1 | SW global and diffuse | ASFS-2.5m | 1 sec | -40 to 70C | 5-15 VDC | <0.1 W |  | up to 20 W | analog/RS232 |  |  | 6600 | ASFS | SPN1/D-BP
Sonic | Metek uSonic-3 Cage | 9 component radial winds, 9 T, 3 cartesian wind components, T, vel, dir | ASFS-3 m; Tower top-11m,30m | max 30 Hz | -40 to 60 C | 10-24 VDC | 2.5 W |  | selectable, 55W, 100W (requires 24 VDC) | RS422, RS485, Ethernet |  | omnidirectional, Ian had ice free with 50W heaters | 9000 | ASFS, Tower | Options: Heat, inclinometer. Do not need any of the other options as we will build our own cables
| Metek uSonic-3 MP | 9 component radial winds, 9 T, 3 cartesian wind components, T, vel, dir | Tower Mid- 2m,5m | max 30 Hz | -40 to 60 C | 10-24 VDC | 2.5 W |  | Selectable, max 100 W | RS422, RS485, Ethernet |  | 20-deg blind zone, Ian had ice free with 50W heaters | 12000 | Tower
Licor | Licor LI-7500DS open path | CO2, H2O mixing ratios | ASFS-2m (summer?): Tower-2m,5m | 5, 10, 20 Hz | -25 to 50 C (possible to -40C) | 10.5 - 30 VDC | 4W typical, 8W max over -25 to 50C. Vendor suggests 6-7W for 5C option. | 30 W | 6-7 W + any additional | Ethernet |  | Housing set poitn at 30C or 5C (later for us!). Need to consider heating mirror element since this is not heated | ~$18500 | ASFS, Tower
| Licor LI-7200RS closed path | CO2, H2O mixing ratios | Tower | 5, 10, 20 Hz | -25 to 50 C (possible to -40C) | 10.5 - 30 VDC | 12 W + <16W for flow module | 30 W | Selectable 0-6 W | Ethernet, RS232 |  | System comes with a controllable heated and insulated inlet. Silicon inlet cap at very end is not heated. | ~$29500 | Tower
Met | Vaisala HMT337 | T, RH |  | 10-sec | -40 to 180 C | 10-35 VDC | 1.5 W |  | 3 W typical, on/off, controlled based on RH | RS-232, RS-485, ethernet |  | T_acc=0.2C (20C) RH_acc=~<3% | 2000 | Tower | Have these in Eureka
| Vaisala WXT535 | P, T, RH, rain | at 2.5m height |  | -52 to 60 C | 6-24 VDC | 0.05 W | 0.09 W | 9.6 W typical, on/off, controlled based on T |  |  | T_acc=0.3C (20C) T_res=0.1C RH_acc=5% RH_res=0.1% | ~$2200 | ASFS, Tower
| Vaisala WXT534 | P, T, RH | at 1m height |  | -52 to 60 C | 6-24 VDC | 0.05 W | 0.09 W | 9.6 W typical, on/off, controlled based on T | SDI-12, RS232, RS485, RS422 |  |  | ~$2100 | ASFS, Tower
Delta-T | Campbell Sci, ASPTC | differential T | From 2-5m, and 2-10m |  | -10 to 70C | 9-13 VDC | 3.12 W | [6 empty] | Tower | Have these in Eureka
| Campbell Sci, FWx | differential T | From 2-5m, and 2-10m |  | ? Recommended for T<-10C |  | 0 W | [6 empty] | Tower
Barometer, high accuracy | Vaisala PTB330 (Class A) | P | ~1m |  | -40 to 60 C | 10-35 VDC | 1 W | 1 W |  | RS232C, RS485, RS422 |  | Accuracy: 0.1 hPa Resolution: 0.01 hPa Response time: 2 sec Stability: 0.1 hPa/yr | 2500 | ASFS
| Paroscientific, Model 6000-16B (6000-16B-IS?) w/ high performance pressure port (1745-103) | P | ~1m | as specified | -54 to 60 C | 6-16 VDC | 0.5 W | 0.5 W |  | RS232, RS485 |  | Accuracy: 0.08 hPa Resolution: 0.001 hPa Stability: 0.1 hPa/yr |  | ASFS
IRT | Apogee SI-421 | Ts | ~2.5m | as specified | -55 to 80 C | 4.5-24 VDC | <0.1W | 0.15 W |  | SDI-12, ModBus |  | 4 different FOV options; also have analog version, accuracy 0.2C | 800 | ASFS, Tower
| Heitronics KT15.85-IIP | Ts | ~2.5m |  | -20 to +60C | 10.5-32 VDC | <3.5W |  |  | RS-232 |  | Accuracy: 0.5 C | 5781 | ASFS, Tower
| Heitronics CT15.85 | Ts | ~2.5m |  | -20 to +60C (can request to -100C) | 10.5-30 VDC | <3.5W |  |  | RS-232, RS-485 |  | Accuracy: 0.5 C | 5346 | ASFS, Tower
Surface height | Campbell Scientific SR50A-316SS ("marine"), or SR50AH (heated) | surface height | ~2.5m | up to <1s. We can sample once every 1-5 minutes | -45 to 50 C | 9-18 VDC | 4.5 W active, 0.02 W passive |  | 3W typical for heater version | SDI-12, RS232, RS485 |  | We should not use heat |  | ASFS, Tower
Sub-surface flux | Hukseflux HFP01SC (flux plate) | surface heat flux | ice-snow interface | variable | -30 to 70 C | 9-15 VDC | 0.02 W (total) |  |  | analog |  | self calibrating | $1015 ($595) | ASFS, Tower
| Hukseflux FHF02 (thermopile, thermocouple) | surface heat flux | ice-snow interface | variable | -40 to 150 C | passive | 0 W |  |  | analog |  |  | 458 | ASFS, Tower
| Campbell Sci, SDI-12 | sub-surface heat flux | through ice and snow |  | -55 to 85 C | 9-28 VDC | <1.4 W |  |  |  |  | min spacing 15 cm max # 36 accuracy 0.2 C |  | ASFS, Tower
| Generic Thermistor string | sub-surface heat flux | through ice and snow | variable | to -40 C | 5-16 V | 0.01 W |  |  | analog |  |  |  | ASFS, Tower
Rain Gauge | OSI OWI-650 | rain rate, rain accumulation | ASFS-2m |  | -40 to 60 C | 10-18 VDC |  | 1.35 W | 7.2 W | RS232 polled
IMU / GPS | [6 empty] | 1.5 W | [6 empty] | ASFS
Camera | [13 empty] | ASFS, Tower
Warning Lights |  |  | ASFS-2.5m; Tower-10m,30m | [10 empty] | ASFS, Tower
Data Logger | Campbell Sci, CR1000X |  |  |  | -40 to 70C | 10-16 VDC | 0.66-1.5 W at 20hZ sampling |  |  | (in) analog, SDI-12, RS-232, RS-485, (out) USB, ethernet |  | onboard microSD storage |  | ASFS, Tower
| MOXA UC-8416/8418 |  |  |  | -40 to 75C | 12-48 VDC | 15 W |  |  | RS-232/422/485, CANbus, USB, ethernet |  | onboard CompactFlash storage (up to many 10s of GB), Runs Linux or Windows CE |  | ASFS, Tower
Link with data logger for comms? | MOXA UC-8100-ME-T |  |  |  | -40 to 85 C | 12-36 VDC | 6 W |  |  | RS-232/422/485, USB, ethernet |  | onboard SDHC/XC (up to 2 TB storage), Runs Linux |  | ASFS