## Slide 1
MOSAiC
Multidisciplinary drifting Observatory
for the Study of Arctic Climate
www.mosaicobservatory.org
Remote Atmospheric Flux Station Planning
Notes: Introduce the organizing team and science plan writing team.

## Slide 2
Remote Atmospheric Flux Stations (RAFS) - Structural Issues
measurement interactions/interference: - snow drift, turbulence generation, flow distortion, radiation impacts, melting/settling (out-of-level)
b) “sled”, box, or “jack stands” – shape, mobility, snow drift
c) box integrated with sonics & instruments, or separate box & tripod? – compactness, transportability, measurement interference
d) transportability, interactions with surface, safety as surface changes (melts) (flotation)
e) serviceability – height, step ladder, stakes/anchors, etc

## Slide 3
radiation
snow depth
met
sonic
barometer
licor
irt
wind generator
solar panel
Sled and/or flotation system
antenna
camera
Atmospheric Surface Flux Station – Conceptual Design

## Slide 4
radiation
snow depth
met
sonic
baro
licor
irt
wind generator
solar panel
Sled and/or flotation system
antenna
camera
Atmospheric Surface Flux Station – Conceptual Design

## Slide 5
Box: E-FOY Fuel Cell power supply/data acquisition

## Slide 6
110 W, 24V.

## Slide 7
NCAR PAM Stations
AOE-2001 –late August 2001 – helicopter deployments
SHEBA
SHEBA
Criticisms from J. Militzer:
- individual sled runners too short, difficult to move
- radiometers/boom too heavy - need for wind turbine uncertain - box/footpads created own meltponds - snow drift around box

## Slide 8
Measured wind speed and momentum flux
distance to ice ridge (m)
height (m)
Measured wind speed (contour lines) and momentum flux (color) downstream of a large (4m) pressure ridge (Garbrecht et al., 1999) – significant variability over a few 100 m
Impacts of Structure on Measurements
Small-Scale Spatial Variability: Ridge impact on the surface layer variables
Also small obstacles > 5-10 cm height strongly influence the atmospheric surface drag!
Important implications for upwind contamination of turbulence measurements!

## Slide 9
Relative Wind Direction (deg)
Ship Structure Impacts on Sampled Airflow - R/V Sikuliaq
Bow-on winds
Starboard winds
Stern winds
Stern winds
Port winds
Rel Wind – Rel Windref (m/s)
- mast port/starboard sensor used as reference for ship-relative winds (SHR winds) from port/starboard - mast starboard/port sensor differs by -3 m/s to +2 m/s of port/starboard sensor for SHR winds from port/starboard - PSD upper bow winds between 0 to +1 m/s of reference wind for SHR winds between ±135° - PSD lower bow winds 0 – 1 m/s lower than PSD upper winds - 1-2 m/s acceleration near ±50-70° - 1-4 m/s deceleration with winds from stern - ship bow 2D sonic (“Ship SCS”) shows assymetric acceleration relative to PSD upper (which it is right next to)
Airflow tilt at PSD Sonics
Impacts also on a) turbulence/fluxes, b) air pressure

## Slide 10
radiation
snow depth
met
sonic
barometer
licor
irt
solar panel
Sled or flotation system
~5° airflow tilt &
5% wind reduction
From side:
~5° airflow tilt &
12%? wind acceleration
With this configuration, need airflow study, or pre-experiment field flow study (~ 2 months); post field-program corrections
Impacts of Structure on Measurements – based on ship experience
2.5 m
1.0 m
Other possible issues: a) larger impact of box on outgoing radiation, b) solar heating of box (heat island) impacting Hs; c) proximity of H20 exhaust from fuel cells impacting humidity(?); d) heavier/more cumbersome to sling as integrated unit with helicopter.

## Slide 11
Suggestion:
Instrument tripod separated from DA box/sled by ~ 5-8 m; minimize sled angle to sonic; remove affected angle from useful data.
Box on ~2-m Nansen sled with or without small duck boat attached on top (500 lbs of flotation max necessity?). Boat needs to be fastened without causing leak. Vendor for Nansen sled? Duck boat? Build Nansen sled?

## Slide 12
Summer meltpond development (Maui) – note reduced melt under ablation shields
Wintertime lead opening (FLA)
Lead closing crushing Cleveland
Ice “disappearance” near Florida
Dangers for RAF Stations
Drastic ice movements (Ice tectonics)
Summer meltponds - 26% (+?) areal coverage
Need for mobility & flotation?

## Slide 13
Nansen Sledge
- Strong for heavy loads
- Sufficiently long to ride easily over rough surface
- High enough and sufficiently sparse supports to reduce drift build-up (?)
SHEBA

## Slide 14
Duck Boat
Flotation in case of melt pond development, lead opening
Duck boat with three 200-lb people plus gear & motor
Duck boat help for this situation?

## Slide 15
Safe pedestal under main tower due to:
a) Wooden posts frozen in with fresh water
b) Plywood & white cloth ablation shields c) luck
SHEBA-late Aug
Meltpond formed by PAM box/sled
Need for step ladder
OTHER STRUCTURAL CONSIDERATIONS
Note pedestal for guy anchor

## Slide 16
Snow Drifting SHEBA PAMs in March

## Slide 17
- tripods in shallow water - pad anchors