| SITE: Tiksi, Russia | DATES OF VISIT: May 22 – June 6, 2012 |
| --- | --- |
| PERSONNEL: Kustov, Bogorodsky |
| Tasking to complete on site (activity, est. time duration of task, materials required): Troubleshoot Rad1 PIR, (4-8 hours), intermittent noise problem in data. If needed, a replacement pyrgeometer dome/thermistor, soldering tools, wire cutters, were sent to Tiksi. RETURN PIR DOME TO NOAA BOULDER AFTER REPAIR OR IF NOT NEEDED 3m Sonic Repair (4-8 hours), replacement circuit board, shim tool, soldering tools, arctic rated cable, connectors, 15m Sonic Reset (2 hours), good weather, laptop with hyperterminal and a serial com port or a usb-serial convertor cable, procedure Sonic Arctic Rate cable fabrication (4-8 hours) , arctic rate cable , connectors for sonic, Remount Licor on tower (2 hours), u-bolts (preferred) or hose clamp (at least 2) Splice tower data cable (4-8 hours), good weather, splice box, connectors, wire strippers/crimpers, wire cutters, tools to strip main data cable (WILL LIKELY WAIT UNTIL AUGUST) Flux Tower RH and Temp sensor troubleshooting CAF Wind Bird troubleshooting (2-4 hours), check connections on tower for the wind bird wires and data logger wires. O3 meter operation/data collection (2-4 hours), install automatic pump switch (if it arrives with Olga), once installed, initiate automatic data downloads using a python script to be supplied by NOAA Ensure proper installation, operation and data collection for the CIMEL CRN maintenance (2-4 hours) perform routine maintenance (see checklist) check aspirator fan operation Report on number of flask boxes which have been shipped to YUGMS/US Shipping & Materials: Items which have been recently shipped to allow completion of this tasking: PIR dome, soldering materials and tools ATI Sonic Board, shim tool, soldering materials, test cables 200 feet of 10 conductor arctic cable, sonic connectors, adapter cables Cable splice box, connectors, tools, mounting hardware (in YUGMS) O3 meter pump switch to be sent with Olga if prepared in time |