| Requested by: Taneil Uttal | Requested to: A. Maksthas, V Kustov |
| --- | --- |
| Date Requested: 5/1/12 | Date Received/Responded to: |
| Priority/critical level (1-5): 3m Sonic Repair (4) 1 = low level; 5 = immediate attention necessary |
| Description of Problem (History, Observations, Analysis, Diagnosis): The 3m Sonic Anemometer stopped collected data in mid-July 2011. It was discovered to have water in the cable during a service trip in September 2011 which is suspected to have shorted out the instrument’s circuit board. Action Items: NOAA Purchase circuit board, cables, connectors, and parts for the sonic and ship to Tiksi. Provide a repair procedure for on site technicians (Provide by M. Okraszewski) TIKSI OBSERVATORY STAFF (assigned to:___________________________) Carry out sonic repair procedure Tracking Date of failure: _~7/16/2011___________________ Date work order sent: __5/1/2012_______________ Date of expected repair: _6/1/2012_______________ Date responded to with understanding of the report: ______________________ Date of action item completion: ______________________ MATERIALS REQUIRED: ATI circuit board and connectors for cables sent by NOAA in box EG537092437US delivered to YUGMS 3/13/12 Soldering materials and tools Test Cable Silicone sealant Flux Computer for testing Reporting (please provide a report on the action items and any problems encountered to matt.okraszewski@noaa.gov and kustov@aari.ru ) |