| Requested by: Taneil Uttal | Requested to: A Maksthas, V Kustov |
| --- | --- |
| Date Requested: 5/4/12 | Date Received/Responded to: |
| Priority/critical level (1-5): 15m Sonic Reset (4) 1 = low level; 5 = immediate attention necessary |
| Description of Problem (History, Observations, Analysis, Diagnosis): Data from the 15m Sonic on the flux tower stopped being collected on Dec 27, 2011. The most probably cause is due to corruption of the instrument settings via a power interruption. Without remote access to the instrument from the CAF (due to the data cable needing repair), the method for fixing the problem is to take a laptop with a serial port and hyperterminal to the tower, plug the instrument cable into computer and carry out the reset procedure. Action Items: NOAA Provide sonic reset procedure (assigned to Matt O) TIKSI OBSERVATORY STAFF (assigned to:___________________________) Carry out sonic reset procedure MATERIALS REQUIRED: Laptop computer containing Hypeterminal software and a serial COM port. If no serial COM port exists, a serial to USB convertor can be used. It is advisable to have a fully charged battery on the laptop as well. Also, if it is a bright sunny day, having a dark cloth or jacket to put over the screen helps too. This is best done on a good weather day with little wind and temperatures above 0C. Tracking Date of failure: _12/27/2011___________________ Date work order sent: __5/1/2012_______________ Date of expected repair: _5/14/2012_______________ Date responded to with understanding of the report: ______________________ Date of action item completion: ______________________ Reporting (please provide a report on the action items and any problems encountered to matt.okraszewski@noaa.gov and kustov@aari.ru ) |