## SCS Sensor List
Type | Manufacturer | Make | Model | SN | SCS Sensor Name | SCS Comm Port | Units | Location | Install Date | Last Calibration | Descriptoin | SCS Comm Port | Sentence Lable | Dat Field Position | Polled Message | Data Type | Log Rate | Baud Rate | Data Bits | Parity | Stop Bits | SCS Folder
NAV | Furono | GP | 150 |  | Primary GPS COG | 11 | Degrees | Antenna on Flying Bridge, Receiver on Bridge |  |  | Primary GPS Course Over Ground | 11 | $GPVTG | 1 |  | NMEA | 1 sec | 4800 | 8 | N | 1 | NAV
NAV | Furono | GP | 150 |  | Primary GPS HDOPS | 11 | # | Antenna on Flying Bridge, Receiver on Bridge |  |  | Primary GPSHDOPS | 11 | $GPGGA | 8 |  | NMEA | 1 sec | 4800 | 8 | N | 1 | NAV
NAV | Furono | GP | 150 |  | Primary GPS QUALITY | 11 | # | Antenna on Flying Bridge, Receiver on Bridge |  |  | Primary GPS Quality | 11 | $GPGGA | 6 |  | NMEA | 1 sec | 4800 | 8 | N | 1 | NAV
NAV | Furono | GP | 150 |  | Primary GPS SATS | 11 | # | Antenna on Flying Bridge, Receiver on Bridge |  |  | Primary GPS Satelites | 11 | $GPGGA | 7 |  | NMEA | 1 sec | 4800 | 8 | N | 1 | NAV
NAV | Furono | GP | 150 |  | Primary GPS SOG | 11 | Knots | Antenna on Flying Bridge, Receiver on Bridge |  |  | Primary GPS Speed Over Ground | 11 | $GPVTG | 5 |  | NMEA | 1 sec | 4800 | 8 | N | 1 | NAV
NAV | Furono | GP | 150 |  | Primary GPS Distance | 11 | Nautical Miles | Antenna on Flying Bridge, Receiver on Bridge |  |  | Primary GPS Distance | 11 | $GPBWR | 10 |  | NMEA | 1 sec | 4800 | 8 | N | 1 | NAV
NAV | Furono | GP | 150 |  | Primary GPS LAT | 11 | Decimal Degrees | Antenna on Flying Bridge, Receiver on Bridge |  |  | Primary GPS Latitude | 11 | $GPGGA | 2 |  | NMEA | 1 sec | 4800 | 8 | N | 1 | NAV
NAV | Furono | GP | 150 |  | Primary GPS LON | 11 | Decimal Degrees | Antenna on Flying Bridge, Receiver on Bridge |  |  | Primary GPS Longitude | 11 | $GPGGA | 4 |  | NMEA | 1 sec | 4800 | 8 | N | 1 | NAV
NAV | Furono | GP | 150 |  | Primary GPS Time | 11 | hhmmss | Antenna on Flying Bridge, Receiver on Bridge |  |  | Primary GPS Time | 11 | $GPGGA | 1 |  | NMEA | 1 sec | 4800 | 8 | N | 1 | NAV
NAV | Furono | GP | 150 |  | Secondary GPS COG | 12 | Degrees | Antenna on Flying Bridge, Receiver on Bridge |  |  | Secondary GPS Course Over Ground | 12 | $GPVTG | 1 |  | NMEA | 1 sec | 4800 | 8 | N | 1 | NAV
NAV | Furono | GP | 150 |  | Secondary GPS HDOPS | 12 | # | Antenna on Flying Bridge, Receiver on Bridge |  |  | Secondary GPS HDOPS | 12 | $GPGGA | 8 |  | NMEA | 1 sec | 4800 | 8 | N | 1 | NAV
NAV | Furono | GP | 150 |  | Secondary GPS LAT | 12 | Decimal Degrees | Antenna on Flying Bridge, Receiver on Bridge |  |  | Secondary GPS Latitude | 12 | $GPGGA | 2 |  | NMEA | 1 sec | 4800 | 8 | N | 1 | NAV
NAV | Furono | GP | 150 |  | Secondary GPS LON | 12 | Decimal Degrees | Antenna on Flying Bridge, Receiver on Bridge |  |  | Secondary GPS Longitude | 12 | $GPGGA | 4 |  | NMEA | 1 sec | 4800 | 8 | N | 1 | NAV
NAV | Furono | GP | 150 |  | Secondary GPS Quality | 12 | # | Antenna on Flying Bridge, Receiver on Bridge |  |  | Secondary GPS Quality | 12 | $GPGGA | 6 |  | NMEA | 1 sec | 4800 | 8 | N | 1 | NAV
NAV | Furono | GP | 150 |  | Secondary GPS SATS | 12 | # | Antenna on Flying Bridge, Receiver on Bridge |  |  | Secondary GPS Satelites | 12 | $GPGGA | 7 |  | NMEA | 1 sec | 4800 | 8 | N | 1 | NAV
NAV | Furono | GP | 150 |  | Secondary GPS SOG | 12 | Knots | Antenna on Flying Bridge, Receiver on Bridge |  |  | Secondary GPS Speed Over Ground | 12 | $GPVTG | 5 |  | NMEA | 1 sec | 4800 | 8 | N | 1 | NAV
NAV | Furono | GP | 150 |  | Secondary GPS Time | 12 | hhmmss | Antenna on Flying Bridge, Receiver on Bridge |  |  | Secondary GPS Time | 12 | $GPGGA | 1 |  | NMEA | 1 sec | 4800 | 8 | N | 1 | NAV
NAV | Teledyne | Meridian Standard |  |  | GYRO Heading | 14 | Degrees | Gyro/IMU room | 2012 |  | Gyro | 14 | $HEHDT | 1 |  | NMEA | 1 Sec | 4800 | 8 | N | 1 | Heading
NAV | TSS |  |  |  | MAHRS-HEADING | 16 | Degrees | Gyro/IMU room |  |  | Microflex Attitude and Heading Reference System | 16 | $HEHDT | 1 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | Heading
NAV | Applanix | POSMV | V5 | 6494 | POSMV Accuracy Hdg | 17 | # | Computer Lab/IMU Room/Flying Bridge | 2010-08-23 |  | POSMV Heading Accuracy | 17 | $INGGA | 9 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | NAV
NAV | Applanix | POSMV | V5 | 6494 | POSMV Accuracy Pitch | 17 | # | Computer Lab/IMU Room/Flying Bridge | 2010-08-23 |  | POSMV Pitch Accuracy | 17 | $INGGA | 8 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | NAV
NAV | Applanix | POSMV | V5 | 6494 | POSMV Accuracy Roll | 17 | # | Computer Lab/IMU Room/Flying Bridge | 2010-08-23 |  | POSMV Roll Accuracy | 17 | $INGGA | 7 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | NAV
NAV | Applanix | POSMV | V5 | 6494 | POSMV GAMS | 17 | # | Computer Lab/IMU Room/Flying Bridge | 2010-08-23 |  | POSMV GAMMS | 17 | $INGGA | 10 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | NAV
NAV | Applanix | POSMV | V5 | 6494 | POSMV Heading | 17 | Degrees | Computer Lab/IMU Room/Flying Bridge | 2010-08-23 |  | POSMV Heading | 17 | $INGGA | 2 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | NAV
NAV | Applanix | POSMV | V5 | 6494 | POSMV Heave | 17 | Meters | Computer Lab/IMU Room/Flying Bridge | 2010-08-23 |  | POSMV Heave | 17 | $INGGA | 6 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | NAV
NAV | Applanix | POSMV | V5 | 6494 | POSMV IMU | 17 | # | Computer Lab/IMU Room/Flying Bridge | 2010-08-23 |  | POSMV IMU | 17 | $INGGA | 11 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | NAV
NAV | Applanix | POSMV | V5 | 6494 | POSMV LAT | 17 | Decimal Degrees | Computer Lab/IMU Room/Flying Bridge | 2010-08-23 |  | POSMV Latitude | 17 | $INGGA | 2 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | NAV
NAV | Applanix | POSMV | V5 | 6494 | POSMV LON | 17 | Decimal Degrees | Computer Lab/IMU Room/Flying Bridge | 2010-08-23 |  | POSMV Longitude | 17 | $INGGA | 4 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | NAV
NAV | Applanix | POSMV | V5 | 6494 | POSMV Pitch | 17 | Degrees | Computer Lab/IMU Room/Flying Bridge | 2010-08-23 |  | POSMV Pitch | 17 | $INGGA | 4 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | NAV
NAV | Applanix | POSMV | V5 | 6494 | POSMV QUALITY | 17 | # | Computer Lab/IMU Room/Flying Bridge | 2010-08-23 |  | POSMV Quality | 17 | $INGGA | 6 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | NAV
NAV | Applanix | POSMV | V5 | 6494 | POSMV Roll | 17 | Degrees | Computer Lab/IMU Room/Flying Bridge | 2010-08-23 |  | POSMV Roll | 17 | $INGGA | 5 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | NAV
NAV | Applanix | POSMV | V5 | 6494 | POSMV SATS | 17 | # | Computer Lab/IMU Room/Flying Bridge | 2010-08-23 |  | POSMV Satelites | 17 | $INGGA | 7 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | NAV
NAV | Applanix | POSMV | V5 | 6494 | POSMV Time | 17 | hhmmss | Computer Lab/IMU Room/Flying Bridge | 2010-08-23 |  | POSMV Time | 17 | $INGGA | 1 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | NAV
NAV | Applanix | POSMV | V5 | 6494 | POSMV Time | 17 | hhmmss | Computer Lab/IMU Room/Flying Bridge | 2010-08-23 |  | POSMV Time | 17 | $INGGA | 1 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | NAV
OCN | SeaBird | TSG | SBE21 | 2531 | TSG- Salinity | 18 | PSU | Bow Thruster Void in SeaChest | 2014-10-21 | 2010-06-01 | Sea Surface Salinity | 18 |  | 37 to 42 |  | Serial | 5 sec | 4800 | 8 | N | 1 | Ocean
OCN | SeaBird | TSG | SBE21 | 2531 | TSG- Water-Temperature | 18 | Deg C | Bow Thruster Void in SeaChest | 2014-10-21 | 2010-06-01 | Sea Surface Temperature | 18 |  | 9 to 14 |  | Serial | 5 sec | 4800 | 8 | N | 1 | Ocean
OCN | SeaBird | TSG | SBE21 | 2531 | TSG-Conductivity | 18 | S/M | Bow Thruster Void in SeaChest | 2014-10-21 | 2010-06-01 | Sea Surface Conductivity | 18 |  | 24 to 28 |  | Serial | 5 sec | 4800 | 8 | N | 1 | Ocean
OCN | SeaBird | TSG | SBE21 | 2531 | TSG-Sound Velocity | 18 | M/Sec | Bow Thruster Void in SeaChest | 2014-10-21 | 2010-06-01 | Sea Surface Sound Velocity | 18 |  | 49 to 54 |  | Serial | 5 sec | 4800 | 8 | N | 1 | Ocean
OCN | Turner | Au10 | 5 | 5249 FRXX | Fluro | 19 | volts | Hydro Lab |  |  | Flurormeter | 19 |  | 31-37 |  | Serial | 1 sec | 9600 | 8 | N | 1 | Ocean
Winch | Markey | DESH5 | FWD |  | Traction Winch Rate | 22 | M/Min | Aft Traction Winch |  |  | Aft Traction Winch Speed | 22 |  | 8 to 10 |  | Serial | 1 sec | 9600 | 8 | N | 1 | Winch
Winch | Markey | DESH5 | FWD |  | Traction Winch Tension | 22 | LBS | Aft Traction Winch |  |  | Aft Traction Winch Tension | 22 |  | 12 to 16 |  | Serial | 1 sec | 9600 | 8 | N | 1 | Winch
Winch | Markey | DESH5 | FWD |  | Traction Winch line Out | 22 | Meters | Aft Traction Winch |  |  | Aft Traction Winch Line Out | 22 |  | 2 to 6 |  | Serial | 1 sec | 9600 | 8 | N | 1 | Winch
Winch | Markey | DESH5 | FWD |  | Hydro Winch Line Out | 23 | Meters | Forward Hydrographic Winch |  |  | FWD Hydro Winch Line Out | 23 | RD | 3 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | Winch
Winch | Markey | DESH5 | FWD |  | Hydro Winch Rate | 23 | M/Min | Forward Hydrographic Winch |  |  | FWD Hydro Winch Line Speed | 23 | RD | 2 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | Winch
Winch | Markey | DESH5 | FWD |  | Hydro Winch Tension | 23 | LBS | Forward Hydrographic Winch |  |  | FWD Hydro Winch Tension | 23 | RD | 1 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | Winch
NAV | SyQuest | Bathy | 2010 | 2611 | Bathy-Depth- Fathoms | 25 | Fathoms | Computer lab/Transducer Void | 2010-08-20 |  | Sub Bottom Profiler and Bathymetric EchoSounder | 25 | $SDDBT | 5 |  | NMEA | 30 sec | 9600 | 8 | N | 1 | Bathy
NAV | SyQuest | Bathy | 2010 | 2611 | Bathy-Depth- Feet | 25 | Feet | Computer lab/Transducer Void | 2010-08-20 |  | Sub Bottom Profiler and Bathymetric EchoSounder | 25 | $SDDBT | 1 |  | NMEA | 30 sec | 9600 | 8 | N | 1 | Bathy
NAV | SyQuest | Bathy | 2010 | 2611 | Bathy-Depth- Meters | 25 | Meters | Computer lab/Transducer Void | 2010-08-20 |  | Sub Bottom Profiler and Bathymetric EchoSounder | 25 | $SDDBT | 3 |  | NMEA | 30 sec | 9600 | 8 | N | 1 | Bathy
MET | Vaisala | PTB | 330 | F36100001 | Baro | 39 | mb | Bridge Radio Room Port Bulkhead |  |  | Barometer Approx. sensor height 15m | 39 |  | 1 to 7 |  | Serial | 1 sec | 9600 | 8 | N | 1 | Met
MET | Vaisala | PTB | 330 | F36100001 | Baro | 39 | mb | Bridge Radio Room Port Bulkhead |  |  | Barometer Value Corrected for Sea Level | 39 |  | 1 |  | Serial | 1 sec | 9600 | 8 | N | 1 | Met
OCN | SeaBird | CTD | 911+ | 483 | SBE11/9 Depth | 40 | Meters | CTD Carousel Staging Bay | Aug. 2014 | 2014-05-07 | CTD SBE9+ Depth Calculated | 40 |  | 7 to 15 |  | Serial | 1 sec | 9600 | 8 | N | 1 | CTD
OCN | SeaBird | CTD | 911+ | 483 | SBE11/9 Pressure | 40 | db | CTD Carousel Staging Bay | Aug. 2014 | 2014-05-07 | CTD SBE9+ Pressure | 40 |  | 22 to 30 |  | Serial | 1 sec | 9600 | 8 | N | 1 | CTD
OCN | SeaBird | CTD | 911+/SBE4C | 1473 | SBE11/9 Salinity | 40 | PSU | CTD Carousel Staging Bay | Aug. 2014 | 2014-05-07 | CTD Salinity from Primary Conductivity SBE 4C sensor SN1473 | 40 |  | 73 to 80 |  | Serial | 1 sec | 9600 | 8 | N | 1 | CTD
OCN | SeaBird | CTD | 911+ | 483 | SBE11/9 Scan | 40 | # | CTD Carousel Staging Bay | Aug. 2014 | 2014-05-07 | CTD SBE9+ Scan Count | 40 |  | 102 to 105 |  | Serial | 1 sec | 9600 | 8 | N | 1 | CTD
OCN | SeaBird | CTD | 911+ | 483 | SBE11/9 Sound Velocity | 40 | M/Sec | CTD Carousel Staging Bay | Aug. 2014 | 2014-05-07 | CTD SBE9+ Sound Velocity calculated | 40 |  | 87 to 94 |  | Serial | 1 sec | 9600 | 8 | N | 1 | CTD
OCN | SeaBird | CTD | 911+/SBE3T | 2866 | SBE11/9 Temp | 40 | Dec C | CTD Carousel Staging Bay | Aug. 2014 | 2014-05-07 | CTD Temperature from Primary Temp SBE 3T sensor SN4866 | 40 |  | 39 to 46 |  | Serial | 1 sec | 9600 | 8 | N | 1 | CTD
OCN | IMET |  |  | ??? | IMET-SST | 43 | Deg C | Inside Hull Mounted in Bow Thruster Void FWD of SeaChest |  |  | Sea Surface Temperature | 43 |  | 1 | #SST01C | Polled | 1 sec | 9600 | 8 | N | 1 | Met | Removed for Calibration
MET | Eppley | PIR |  | LWR301 | IMET-LWR | 44 | Wats/m2 | Fwd Deck in front of Bridge |  |  | Long Wave Radiation | 44 |  | 26 to 30 | #LWR01C | Polled | 1 sec | 9600 | 8 | N | 1 | Met
MET | Eppley | PSP |  | SWR258 | IMET SWR | 45 | Wats/m2 | Fwd Deck in front of Bridge | 2015-01-14 |  | Short Wave Radiation | 45 |  | 11 to 17 | #SWR01C | Polled | 1 sec | 9600 | 8 | N | 1 | Met
MET | RM Young |  | 5103 | WND242 | IMET -Rwind-Dir | 47 | Degrees | Fwd Bow Mast 14.1m above sea level |  |  | Relative Wind Direction | 47 |  | 12 to 17 | #WND01A | Serial | 1 sec | 9600 | 8 | N | 1 | Met
MET | RM Young |  | 5103 | WND242 | IMET -Rwind-Spd-Msecs | 47 | M/Sec | Fwd Bow Mast 14.1m above sea level |  |  | Relative Wind Speed | 47 |  | 7 to 10 | #WMD01A | Serial | 1 sec | 9600 | 8 | N | 1 | Met
MET | Rotronics | MP | 101A | HRH??? | IMET-Air Temp | 49 | Degrees | FWD Bow Mast 12.98m above sea level |  |  | Humidity/Air Temp Sensor | 49 |  | 11 to 17 | #HRH01C | Polled | 1 Sec | 9600 | 8 | N | 1 | Met
MET | Rotronics | MP | 101A | HRH??? | IMET-Relative Humidity | 49 | Percent | FWD Bow Mast 12.98m above sea level |  |  | Humidity/Air Temp Sensor | 49 |  | 2 to 8 | #HRH01C | Polled | 1 Sec | 9600 | 8 | N | 1 | Met
MET | RMYoung |  | 50201 | PRC251 | IMET-Rain | 50 | Milimeters | FWD Bow Mast 13.99m above sea level |  |  | Precipitation Sensor | 50 |  | 3 to 8 | #PRC01C | Polled | 1 Sec | 9600 | 8 | N | 1 | Rain | Removed duirng RB1501 due to sensor failure
OCN | Reson | SVP | 70 |  | SVP-Port-sound velocity | 53 | m/sec | Port Hull |  | 2014-04-06 | Port Sound Velocity Sensor | 53 |  | 2 to 8 |  | Serial | 1 sec | 9600 | 8 | N | 1 | Ocean
OCN | Reson | SVP | 70 |  | SVP-STBD-sound velocity | 54 | m/sec | STBD Hull |  | 2014-04-06 | STBD Sound Velocity Sensor | 54 |  | 2 to 8 |  | Serial | 1 sec | 9600 | 8 | N | 1 | Ocean
MET | Vaisala |  |  |  | Vaisala WndPort-WIMWV-Bearing | 55 | Degrees | Flying Bridge Port Side |  |  | Port Relative Wind Direction | 55 | $WIMWV | 1 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | Met
MET | Vaisala |  |  |  | Vaisala WndPort-WIMWV-spd | 55 | Knots | Flying Bridge Port Side |  |  | Port Relative Wind Speed | 55 | $WIMWV | 1 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | Met
MET | Vaisala |  |  |  | Vaisala WndStbd-WIMWV-Bearing | 56 | Degrees | Flying Bridge Port Side |  |  | STBD Relative Wind Direction | 56 | $WIMWV | 1 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | Met
MET | Vaisala |  |  |  | Vaisala WndStbd-WIMWV-spd | 56 | Knots | Flying Bridge Port Side |  |  | STBD Relative Wind Direction | 56 | $WIMWV | 1 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | Met
NAV | Kongsburg | EM | 122 | 309652 | Multibeam-Depth-Meters | 58 | Meters | Computer Lab/ Transducer Void |  |  | Kongburg EM122 Deepwater Multibeam | 58 | $KIDPT | 1 |  | NMEA | variable | 9600 | 8 | N | 1 | Multibeam
NAV | RDI Teledyne | OS | 75 | 1558 | ADCP-BinsUsed | 10512 | # | Computer lab/Transducer Void | Transducer Replaced 9/30/2014 |  | Acoustic Doppler Profiler Bins Used | 10512 | $os75bb | 5 |  | TCP | 1 sec |  |  |  |  | ADCP
NAV | RDI Teledyne | OS | 75 | 1558 | ADCP-Dir | 10512 | Degrees | Computer lab/Transducer Void | Transducer Replaced 9/30/2014 |  | Acoustic Doppler Profiler Direction | 10512 | $os75bb | 3 |  | TCP | 1 sec |  |  |  |  | ADCP
NAV | RDI Teledyne | OS | 75 | 1558 | ADCP-Spd | 10512 | Knots | Computer lab/Transducer Void | Transducer Replaced 9/30/2014 |  | Acoustic Doppler Profiler Speed | 10512 | $os75bb | 1 |  | TCP | 1 sec |  |  |  |  | ADCP
NAV | RDI Teledyne | OS | 75 | 1558 | ADC-Time | 10512 | hhmmss | Computer lab/Transducer Void | Transducer Replaced 9/30/2014 |  | Acoustic Doppler Profiler Time | 10512 | $os75bb | 6 |  | TCP | 1 sec |  |  |  |  | ADCP
NAV | Trimble | SPS | 315 | 5205D53008 | Trimble1_COG | 28001 | Degrees | Bridge Chart Table Rack | Shipyard 2012 |  | Trimble GPS Course Over Ground | 28001 | $GPVTG | 1 |  | TCP | 1 sec |  |  |  |  | NAV
NAV | Trimble | SPS | 315 | 5205D53008 | Trimble1_SOG | 28001 | Knots | Bridge Chart Table Rack | Shipyard 2012 |  | Trimble GPS Speed Over Ground | 28001 | $GPVTG | 5 |  | TCP | 1 sec |  |  |  |  | NAV
NAV | Trimble |  |  |  | Trimble2_HDOPS | 28001 | # | Bridge Chart Table Rack | Shipyard 2012 |  | Trimble GPS HDOPS | 28001 | $GPGGA | 5 |  | TCP | 1 sec |  |  |  |  | NAV
NAV | Trimble | SPS | 315 | 5205D53008 | Trimble2_LAT | 28001 | Decimal Degrees | Bridge Chart Table Rack | Shipyard 2012 |  | Trimble GPS Latitude | 28001 | $GPGGA | 5 |  | TCP | 1 sec |  |  |  |  | NAV
NAV | Trimble | SPS | 315 | 5205D53008 | Trimble2_LON | 28001 | Decimal Degrees | Bridge Chart Table Rack | Shipyard 2012 |  | Trimble GPS Longitude | 28001 | $GPGGA | 5 |  | TCP | 1 sec |  |  |  |  | NAV
NAV | Trimble | SPS | 315 | 5205D53008 | Trimble2_Quality | 28001 | # | Bridge Chart Table Rack | Shipyard 2012 |  | Trimble Quality | 28001 | $GPGGA | 5 |  | TCP | 1 sec |  |  |  |  | NAV
NAV | Trimble | SPS | 315 | 5205D53008 | Trimble2_SATS | 28001 | # | Bridge Chart Table Rack | Shipyard 2012 |  | Trimble GPS Satelites | 28001 | $GPGGA | 5 |  | TCP | 1 sec |  |  |  |  | NAV
NAV | Trimble | SPS | 315 | 5205D53008 | Trimble2_Time | 28001 | hhmmss | Bridge Chart Table Rack | Shipyard 2012 |  | Trimble GPS Time | 28001 | $GPGGA | 5 |  | TCP | 1 sec |  |  |  |  | NAV
NAV | Furono | GP | 150 |  | GP150GPS-SOG-Msec-Value |  | M/Sec | Antenna on Flying Bridge, Receiver on Bridge |  |  | Derived Sensor from Primary GPS |  | $DERV | 1 |  | Derived | 1 sec |  |  |  |  | NAV
MET | RM Young |  | 5103 | WND301 | IMET Rwind-Spd-Knts-VALUE |  | Knots | Fwd Bow Mast 14.1m above sea level |  |  | Relative Wind Speed |  | $DERV | 1 |  | Derived | 1 sec | 9600 | 8 | N | 1 | Met
MET | RM Young |  | 5103 | WND301 | IMET Twind-SPEED |  | Knots | Fwd Bow Mast 14.1m above sea level |  |  | True Wind Speed |  | $DERV | 1 |  | Derived | 1 sec | 9600 | 8 | N | 1 | Met
MET | RM Young |  | 5103 | WND301 | IMET Twind-SPEED |  | M/Sec | Fwd Bow Mast 14.1m above sea level |  |  | True Wind Speed |  | $DERV | 1 |  | Derived | 1 sec | 9600 | 8 | N | 1 | Met
MET | RM Young |  | 5103 | WND301 | IMET-Rwind-DirAvg-VALUE |  | Degrees | Fwd Bow Mast 14.1m above sea level |  |  | Relative Wind Direction Average |  | $DERV | 1 |  | Derived | 1 Sec | 9600 | 8 | N | 1 | Met
MET | RM Young |  | 5103 | WND301 | IMET-Twind-DIRECTION |  | Degrees | Fwd Bow Mast 14.1m above sea level |  |  | True Wind Direction Average |  | $DERV | 1 |  | Derived | 1 Sec | 9600 | 8 | N | 1 | Met
MET | Vaisala |  |  |  | Port- Twind-DIRECTION |  | Degrees | Flying Bridge Port Side | 2012-06-11 |  | Derived True Wind Direction from Port Ultrasonic Wind Sensor |  | $DERV | 2 |  | Derived | 1 sec |  |  |  |  | Met
MET | Vaisala |  |  |  | Port- Twind-SPEED |  | Knots | Flying Bridge Port Side | 2012-06-11 |  | Derived True Wind Speed from Port Ultrasonic Wind Sensor |  | $DERV | 1 |  | Derived | 1 sec |  |  |  |  | Met
MET | Vaisala |  |  |  | Stbd- Twind-DIRECTION |  | Degrees | Flying Bridge Port Side | 2012-06-11 |  | Derived True Wind Direction from STBD Ultrasonic Wind Sensor |  | $DERV | 2 |  | Derived | 1 sec |  |  |  |  | Met
MET | Vaisala |  |  |  | Stbd- Twind-SPEED |  | Knots | Flying Bridge Port Side | 2012-06-11 |  | Derived True Wind Speed from STBD Ultrasonic Wind Sensor |  | $DERV | 1 |  | Derived | 1 sec |  |  |  |  | Met
MET | Optical Scientific Inc |  |  |  | precip5-port03 | 41 | mm/hr | 03 deck port side | 2015-01-29 |  | Optical Rain Gauge | 41 | $WIMRA | 5 |  | NMEA | 1 sec | 9600 | 8 | N | 1 | Met | Sensor installed by scientific party in San Francisco on 1/29/2015

## NMEA Sentence Labels
Manufacturer | Make | Model | Description | Data Type | Sentence Label | Field Pos. 1 | Field Pos. 2 | Field Pos. 3 | Field Pos. 4 | Field Pos. 5 | Field Pos. 6 | Field Pos. 7 | Field Pos. 8 | Field Pos. 9 | Field Pos. 10 | Field Pos. 11 | Field Pos. 12 | Field Pos. 13
SyQuest | Bathy | 2010 | Sub Bottom Profiler and Bathymetric EchoSounder | NMEA | $SDDBT | Depth Below Transducer- Feet | f | Depth Below Transducer - Meters | m | Depth Below Transducer- Fathoms | F
Furono | GP | 150 | Primary and Secondary GPS | NMEA | $GPBWR | Time UTC of Observation | Waypoint Latitude | Waypoint Longitude | Bearing Degrees True | Bearing, degrees Magnetic | Distance Nautical Miles | Waypoint ID | Mode Indicator
Furono | GP | 150 | Primary and Secondary GPS | NMEA | $GPGGA | Time UTC of Position | Latitude | Longitude | GPS Quality 0=no fix, 1=GPS, 2=dGPS | # of Satelites in use, 0-12 | Horizontal Dillution of Positin (HDOP) | Antenna Altitude above mean sea level | Geoidal Speration, m | Age of dGPS data | Differential Reference Staiton ID
Teledyne | Meridian Standard |  | Gyro | NMEA | $HEHDT | True Heading (deg) | TRUE
TSS |  | MAHRS | Microflex Attitude and Heading Reference System | NMEA | $HEHDT
Applanix | POSMV | v5 | POSMV | NMEA | $INGGA | Time (hhhmmss.ss) | Latitude (degMin decimal mins) | North or South | Longitude (degMin decimal mins) | East or West | PGS Quality 0=no fix, 1=C/A Standard, 2 DGPS, 3 GPS PSS Mode, 4=RTK fixed, 5= RTK float, 6=Free interval | # of satelites 0-32 | HDOP | Altitude above mean Sea level | Meters | Null | Null | Age of Differential corrections (sec)
Applanix | POSMV | v5 | POSMV | NMEA | $PASHR | Time (hhhmmss.ss) | True Vessle heading (deg) | TRUE | Roll (-90 to +90 degrees) | Pitch (-90 to +90 degrees) | Heave (-99 to +99 meters) | Accuracy Roll (0-9.99 degrees) | Accuracy Pitch (0 to 9.99 degrees) | Accuracy Heading (0 to 9.99 degrees) | Flag Accuracy (0=no aiding, 1 = GPS aiding, 2 = GPS & GMAS aiding | Flag - IMUI (0= IMU Out, 1 = Satisfactory)
SyQuest | Bathy | 2000 | 3.5 or 12kHz | NMEA | $SDDBT | Depth Below Transducer- Feet | f | Depth Below Transducer - Meters | M | Depth Below Transducer- Fathoms | F
Knudsen | CHIRP | 3260 | Sub Bottom Profiler | NMEA | $SDDBT | Depth Below Transducer- Feet | f | Depth Below Transducer - Meters | M | Depth Below Transducer- Fathoms | F | F
TSS |  |  | Microflex Attitude and Heading Reference System | NMEA | $HEHDT | True Heading (deg) | TRUE
Markee | DESH5 |  | Hydrographic Winch | NMEA | RD | Hydro Winch Tension | Hydro Winch Rate | Hydro Winch Payout
EM | 122 | Multibeam Echosounder | Multibeam - Depth | NMEA | $KIDPT | Multibeam Depth (meters)
SeaBird | SBE | 9+ CTD | SBE 9+ CTD | NMEA
Teledyne | OS | 75kHz | Acoustic Doppler Current Profiler | TCP Port 10512 | $os75bb | ADCP Speed (kts) | KTS | ADCP Directin (Deg) |  | ADCP Bins Used | ADCP Time (HHMMSS)
Trimble1 | SPS | 315 | GPS | TCP Port 28001 | $GPVTG | GPS Course Over Ground (Deg) | T |  | M | Speed Over Ground (kts)
Trimble1 | SPS | 315 | GPS | NMEA | $GPGGA | Time (hhmmss) | Latitude (degMin decimal mins) | North or South | Longitude (degMin decimal mins) | East or West | Quality | # of satelites | HDOP
Trimble2 | SPS | 315 | GPS | TCP Port 28001 | $GPVTG | GPS Course Over Ground (Deg) | T |  | M | Speed Over Ground (kts)
Trimble2 | SPS | 315 | GPS | NMEA | $GPGGA | Time (hhmmss) | Latitude (degMin decimal mins) | North or South | Longitude (degMin decimal mins) | East or West | Quality | # of satelites | HDOP
Vaisala |  |  | Ultrasonic Wind Port | NMEA | $WIMWV | Relative Bearing (degrees) | R | Relative Wind Speed | N
Vaisala |  |  | Ultrasonic Wind STBD | NMEA | $WIMWV | Relative Bearing (degrees) | R | Relative Wind Speed | N
Optical Scientific Inc |  |  | Optical Rain Gauge | NMEA | $WIMRA | disconnected rain gauge | disconnected rain gauge | disconnected rain gauge | disconnected rain gauge | Optical Rain Gauge | disconnected rain gauge

## Serial-Polled Sentence Labels
Manufacturer | Make | Model | Description | Sensor Name | Data Type | Sentence Label | Data | Starting/Ending Characters | Data | Starting/Ending Characters | Data | Starting/Ending Characters | Data | Starting/Ending Characters | Data | Starting/Ending Characters | Data | Starting/Ending Characters
Rotronics | MP | 101A | Relative Humidity/Aire Temp | IMET Humidity - Temp | Polled | #HRH01C | Relative Humidity (%) | 2 to 8 | Air Temp. (Deg. C) | 11 to 17
Eppley | PSP |  | Short Wave Radiometry | IMET - Shortwave Radiation | Polled | #SWR01C | Short Wave Radiation (Watts/m2) | 3 to 8
Eppley | PIR |  | Long Wave Radiometry | IMET LWR | Polled | #LWR01C | Long Wave Radiation (Watts/m2)
RM Young |  | 5103 | Wind Aeronometer | IMET- Rel Wind | Serial |  | Count | 1 to 4 | Rel Wind Spd (m/sec) | 7 to 10 | Rel. Wind Direction (deg) | 12 to 17
IMET |  |  | Sea Surface Temp | IMET SST | Polled | #SST01C | Sea Surface Temprature (Deg C) | 2 to 8
RM Young | IMET |  | IMET Rain Gauge | IMET Rain | Polled | #PRC01C | Rain (mm) | 11 to 15
SeaBird | SBE | 21 | Thermosalinograph | TSG | Serial |  | TSG Water Temp (Deg C) | 9 to 14 | TSG Conductivity S/M) | 24 to 28 | TSG Salinity (PSU) | 37 to 42 | TSG Sound Velocity (M/Sec) | 49 to 54
Turner |  | 10AU | Fluorometer | Fluro | Serial |  | Date Time | 12 to 29 | Raw Voltage (V) | 31 to 37
Viasala | PTB | 330 | Barometer | Barometer | Serial |  | Barometric Pressure (mb) | 1 to 7
Markey | DEH5 |  | Traction Winch | Winch-Traction | Serial |  | Winch Traction Payout (meters) | 2 to 6 | Winch Traction Rate (m/sec) | 8 to 10 | Winch Traction Tension (Lbs)
SeaBird | SBE | 9+ CTD | SBE 9+ CTD | CTD | Serial |  | Depth (meters) | 7 to 15 | Pressure (db) | 22 to 30 | Temp. Deg. C) | 39 to 46 | Salinity (PSU) | 73 to 80 | Sound Velocity | 87 to 94 | Scan Count | 102 to 105
Reson | SVP | 70 | Sound Velocity Port | SVP Port | Serial |  | Sound Velocity (M/Sec) | 2 to 8
Reson | SVP | 70 | Sound Velocity STBD | SVP STBD | Serial |  | Sound Velocity (M/Sec) | 2 to 8