## CAPRICORN_Aerosols
Date |  | 2016-03-13 | 2016-03-14 | 2016-03-15 | 2016-03-16 | 2016-03-17 | 2016-03-18 | 2016-03-19 | 2016-03-20 | 2016-03-21 | 2016-03-22 | 2016-03-23 | 2016-03-24 | 2016-03-25 | 2016-03-26 | 2016-03-27 | 2016-03-28 | 2016-03-29 | 2016-03-30 | 2016-03-31 | 2016-04-01 | 2016-04-02 | 2016-04-03 | 2016-04-04 | 2016-04-05 | 2016-04-06 | 2016-04-07 | 2016-04-08 | 2016-04-09 | 2016-04-10 | 2016-04-11 | 2016-04-12 | 2016-04-13 | 2016-04-14 | 2016-04-15
Operator |  | RH | KL | KL |  | KL | KL | RH/KL | KL | KL | RH | KL | KL | KL | RH | KL | KL | RH | KL | RH | RH | RH | RH | RH | RH | KL | KL/RH | KL | RH | KL/RH | RH
APS | 150.229.233.205
Time (station) | Our computer time (GMT+10), Kaitlyn's was +11 until the 3rd April |  |  |  |  | 10:09AM | 09:23AM | 11:24AM | 13:49 | 09:16AM | 11:47 | 11:37:00AM | 13:27 | 11:45AM | 10:56 | 12:26PM | 13:25 | 13:04 | 13:41 | 16:20 | 11:53 | 10:55 | 12:55 | 12:20 | 12:56 | 10:38AM | 13:03 | 14:46 | 13:35 | 14:53 | 13:01
Date (PC) |  | ? | ? | ? |  | 2016-03-16 | 2016-03-17 | 2016-03-19 | 2016-03-20 | 2016-03-20 | 2016-03-22 | 2016-03-23 | 2016-03-24 | 2016-03-25 | 2016-03-26 | 2016-03-27 | 28/0316 | 2016-03-29 | 2016-03-30 | 2016-03-31 | 2016-04-01 | 2016-04-02 | 2016-04-03 | 2016-04-04 | 2016-04-05 | 2015-04-06 | 2016-04-07 | 2016-04-08 | 2016-04-09 | 2016-04-10 | 2016-04-11
Time (PC) | If this is different to station time, please synchronise clocks - should be set to sync with internet time. | 22:05 | 14:54 | 9:51 |  | 11:08PM | 22:24 | 12:23AM | 2:50AM | 22:15 | 1:47 | 0:36 | 2:27AM | 0:45AM | 0:56 | 1:26AM | 2:26AM | 3:05 | 2:40AM | 6:20 | 1:53 | 0:55 | 2:55 | 2:20 | 2:56 | 0:37AM | 3:02 | 4:45 | 3:35 | 14:51 | 3:01
Check data backed up on ship servers | bar: 2016-04-03 to 2016-04-11
Sample distribution OK? |  | Y | Y | Y |  | Y | Y | Y | Y | A bit noisy | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y
Sheath flow (LPM) | In display menu | 4 | 4 | 4 |  | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4
Inlet flow (LPM) | In display menu | 1 | 1 | 1 | bar: 2016-03-16 to 2016-04-11
Total flow (LPM) | Should be the sum of sheath and inlet | 5 | 5 | 5 |  | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
Laser power (%) | In display menu | 75 | 75 | 75 |  | 75 | 75 | 75 | 75 | 75 | 75 | 75 | 75 | 75 | 75 | 75 | 75 | 75 | 75 | 75 | 75 | 75 | 75 | 75 | 75 | 75 | 75 | 75 | 75 | 75 | 75
Laser current (mA) |  | 52.4 | 53.2 | 52 |  | 52.3 | 51.6 | 53.2 | 52.6 | 52.5 | 53.2 | 52.9 | 52.8 | 53.1 | 52.8 | 52.3 | 48.6 | 48.7 | 48.5 | 49.3 | 49 | 48.8 | 48.8 | 49 | 48 | 49.9 | 51.6 | 51.2 | 51.2 | 50.9 | 51.3
Comments |  | Very low concentrations, misty/foggy day with rain this morning |  |  |  | 2016-03-19: Low concentrations; 2016-04-05: Water droplets on APS and flow aren't moving around much as they usually do
CCNC
Time (station) | Our computer time (GMT+10) | [8 empty] | 10:33 | 11:50 |  | 13:31 | 11:46 | 10:57 | 12:28:00PM | 13:27PM | 12:43 | 13:42 | 16:21 | 11:55 | 10:56 | 12:56 |  | 12:59 | 10:40AM |  | 14:47 | 13:37 | 14:48 | 13:10
Date (PC) |  | ? | ? | ? | [5 empty] | 2016-03-21 | 2016-03-22 |  | 2016-03-24 | 2016-02-25 | 2016-03-26 | 2016-03-27 | 2016-03-28 | 2016-03-29 | 2016-03-30 | 2016-03-31 | 2016-04-01 | 2016-04-02 | 2016-04-03 |  | 2016-04-05 | 2016-04-06 | 2016-04-07 | 2016-04-08 | 2016-04-09 | 2016-04-10 | 2016-04-11
Time (CCNC PC) | If this is different to station time, please synchronise clocks - should be set to sync with internet time. | 22:11 | 14:58 | 10:49 | [5 empty] | 0:33 | 1:50 |  | 2:31AM | 0:46AM | 0:57 | 1:28AM | 2:27AM | 2:44 | 2:42AM | 6:21 | 1:55 | 0:56 | 2:56 |  | 2:59 | 0:39AM |  | 4:46 | 3:37 | 4:48 | 3:10
Check data backed up on ship servers | [5 empty] | 2016-03-29: Yes; 2016-04-03: Y; 2016-04-05: Y; 2015-04-06: Y; 2016-04-07: Y; 2016-04-08: Y; 2016-04-09: Y; 2016-04-10: Y; 2016-04-11: Y
Status box green? (Y/N) | If "NO", and it is not an obvious fix, please email Ruhi.Humphries@csiro.au | Y | Y | Y | bar: 2016-03-20 to 2016-03-22; 2016-03-24 to 2016-04-03; 2016-04-05 to 2016-04-11
Size distribution Gaussian? If not, please describe |  | Yes, but not a great gaussian, likely due to the low concentrations | Yes, but not a great gaussian, likely due to the low concentrations | Y | [5 empty] | Y | Y |  | Y | Y | Y | Y | Y | No, VERY low concentrations though. | Y | Y | Y | No, low concentrations though | Y |  | Y | Y | Y | Y | Y | N | Y
SS tab: current SS |  | 0.5 | 0.5 | 0.5 | [5 empty] | 0.5 | 0.5 |  | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 |  | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
Flow tab: Flow ratio (should be between 10.0 ± 0.1) | If not within the range, it may be a small adjustment in the sheath valve. If not, a flow calibration may be necessary. Please consult manual and inform Ruhi. | Y | Y | Y | [5 empty] | Y | Y |  | 9.5 | 9.59 | 9.5 | 9.53 | 9.57 | 10 | 9.91 | 10 | 10 | 10 | 10 |  | 10 (changed down from 10.5) | 10.15 | 10 | 10.1 | 10 | 10 | 10
OPC tab: Concentration |  | 20 | 44 | 173 | [5 empty] | 30 | 70 |  | 167 | 120 | 180 | 70 | 55 | 15 | 120 | 40 | 320 | 20 | 50 |  | 350 | 66 | 100 | 92 | 50 | zero - Column dried | 50
CCNC water reservoir (mL) | The CCNC column must NOT dry out. Please ensure there is enough milli-Q water present until next check. | Full | 2/3 Full | Full | [5 empty] | 2/3 full | 1/2 full |  | 2/3 full | 1/2 full | OK | 2/3 full | 1/2 full | full | 2/3 full | 200 mL | almost full | 350 ml | 230 ml |  | 900 ml | 900mL | 700 | 600 | 550 | 1000 | 900
Water refilled and drain emptied? (Y/N) | Once refilled, ensure all fittings are secured tightly as this is the most common location for leaks to be introduced. Ensure blue cap and white fittings in each cap are tight. Only use Milli-Q water, and ensure no dirt enters the water during refilling. Drain bottles can be emptied down the sink. | N | N | Y | [5 empty] | N | Y |  | N | Y | N | N | N | Y | N | Y | N | N | N |  | N | N | N | N | N | Y, see comments | N
OPC water trap (comments) | Check OPC water trap. If more than 1 cm deep of liquid is present, empty the bottle. | OK | OK | OK | bar: 2016-03-20 to 2016-03-22; 2016-03-24 to 2016-04-03; 2016-04-05 to 2016-04-11
Check bottom of case for water leakage |  | OK | OK | OK | bar: 2016-03-20 to 2016-03-22; 2016-03-24 to 2016-04-03; 2016-04-05 to 2016-04-11
Zero test concentration (should be <5/sec) | Check after changing the water |  |  | Y | bar: 2016-03-20; 2016-03-26; 2016-03-31 to 2016-04-01; 2016-04-05; 2016-04-11
Comments | [5 empty] | 2016-03-29: Measuring some pollution at present with spikes up to 120; 2016-03-31: See log; 2016-04-01: Issue overnight, see log; 2016-04-10: Discovered column had dried out because the new hose with the new bottle was sticking out of the water. I now have cut the new tubing so that it is the correct length, and also refilled it at the same time.
WEEKLY | bar: 2016-03-22
Measure inlet flow using external flowmeter | Inlet flow should be 500 ccm (0.5 LPM) |  |  | Tested at 11:13pm UTC at 0.494 LPM |  | 2016-03-26: 508.4, tested at 00:32-00:34; 2016-03-31: 508.9; 2016-04-10: 500.9
TSI long SMPS
Time (station) | Our computer time (GMT+10) |  |  |  |  | 10:35 | 09:27AM | 11:27AM | 13:52 | 09:17AM | 11:58 | 11:39AM | 13:28 | 11:52AM | 10:45 | 12:30PM | 13:30 | 12:47 | 13:46 | 16:32 | 11:36 | 10:57 | 12:58 | 12:32 | 13:25 | 10:42AM | 12:57 | 14:49 | 13:40 | 15:02 | 13:12
Date (PC) |  |  |  | ? |  | 2016-03-16 | 2016-03-17 | 2016-03-19 | 2016-03-20 | 2016-03-20 | 2016-03-22 | 2016-03-23 | 2016-03-24 | 2016-03-25 | 2016-03-26 | 2016-03-27 | 2016-03-28 | 2016-03-29 | 2016-03-30 | 2016-03-31 | 2016-04-01 | 2016-04-02 | 2016-04-03 | 2016-04-04 | 2016-04-05 | 2016-04-06 | 2016-04-07 | 2016-04-08 | 2016-04-09 | 2016-04-10 | 2016-04-11
Time (PC) | If this is different to station time, please synchronise clocks - should be set to sync with internet time. |  |  | 0:24 |  | 23:35 | 22:26 | 12:27AM | 2:52AM | 22:17 | 1:58 | 0:38AM | 2:28AM | 0:52AM | 0:45 | 2:31AM | 3:29AM | 3:50 | 2:46AM | 6:32 | 1:36 | 0:57 | 2:58 | 2:32 | 3:25 | 0:42AM | 2:56 | 4:48 | 3:40 | 5:02 | 3:12
Check data backed up on ship servers | bar: 2016-04-03 to 2016-04-11
Sample # | [6 empty] | 207 | 516 | 833 | 1066 | 1389 | 1660 | 1970 | 2239 | 2526 | 2823 | 3122 | 3415 | 285 | 618 | 847 | 1127 | 1439 | 1722 | 2021 | 2276 | 2591 | 2901 | 3176 | 3480 | 3746
Sample distribution OK? | Please describe if not normal |  |  |  |  | OK | OK | OK | OK | Quite noisy | OK | OK | OK | OK | OK | OK | OK | Noisy, due to low concentrations and random sampling of exhaust | OK | OK now, had issues overnight, see log | OK | OK - strong 20 nm mode, with nothing much else | Messy and flat, but consistent with GRIMM | OK | OK | OK | OK | OK | OK | OK | Showing background correction, likely due to low overall concentrations
Zero counts on CPC between scans? | [5 empty] | 0.06 |  | 0 | 0 | 0 | 0 | 0 | 0.01 | 0 | 0 | 0 |  |  | 0.7 |  | Yes | Y | Y |  |  | Y | Y | Y |  |  | Y
Scan control on Analog and size is incrementing | DMA screen | bar: 2016-03-31 to 2016-04-11
SMPS CPC status light green | bar: 2016-03-16 to 2016-04-04; 2016-04-06 to 2016-04-11
CPC butanol reservoir level (OK or filled)? | [5 empty] | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK now, had issues overnight, see log | OK | OK | OK | OK | OK | OK | OK | OK | Refill tomorrow | Refilled | OK
3776 CPC pressure A (~100) | On CPC screen |  |  | 101.5 |  | 100.4 | 100.3 | 101.5 | 101.5 | 101.5 | 101.5 | 101.5 | 101.5 | 101.5 | 101.5 | 101.5 | 101.5 | 101.5 | 101.5 | 101.7 | 99.1 | 99.5 | 101.2 | 100.5 | 98.6 | 100.2 | 99.8 | 100.2 | 98.8 | 98.5 | 100.8
3776 CPC pressure O (~70) | On CPC screen |  |  | 79.6 |  | 77.5 | 77.4 | 79 | 79.5 | 79.7 | 79 | 75.8 | 79.1 | 78.6 | 78.7 | 78.8. | 78.8 | 79 | 78.8 | 77.7 | 77.5 | 76.5 | 78.2 | 77.7 | 76.2 | 77.5 | 77 | 77.5 | 76.2 | 76 | 77.9
3776 CPC pressure N (~3) | On CPC screen |  |  | 3.1 |  | 3 | 3 | 3.1 | 3.1 | 3.1 | 3.1 | 3.1 | 3.1 | 3.1 | 3.1 | 3.1 | 3.1 | 3.1 | 3.1 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3
Flows: Sheath inline (3 LPM) | Check on screen | [6 empty] | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3
Flows: Monodisperse inline | Every few days (1L before, 400mL after the 18th) |  |  |  |  | 2016-03-16: 990.4 (23:42); 2016-03-20: ?? (23:11), 400 (0:35); 2016-03-22: 390; 2016-03-23: 400 (0:43); 2016-03-27: 403 (2:35); 2016-04-03: 409.2
Flows: polydisperse in | Every few days (1L before, 400mL after the 18th) |  |  |  |  | 2016-03-16: 986.8 (23:45); 2016-03-20: 439 (23:06), 400 (0:35); 2016-03-22: 400; 2016-03-23: 433 (0:45); 2016-03-27: 445 (2:37); 2016-04-03: 448.8
Flows: Inlet flow | Every few days (1L before, 400mL after the 18th) |  |  |  |  | 2016-03-26: 1074; 2016-04-03: 456.3; 2016-04-07: 400; 2016-04-10: 404.8
program restarted? | only do if something goes wrong |  |  |  |  | No | No | No | No | No | No | No | No | No | No | No | No | YES | No |  | No | No | No | No | No | No | No | No | No | No | No
New file name (if program restarted) | [5 empty] | 2016-03-29: CAPRICORN_SMPS_TSI_05_UTC
Comments | See text log as well |  |  |  |  | 2016-03-22: Stopped acquisition to check settings; 2016-03-29: Discovered computer time was set to GMT (London), rather than GMT UTC, and daylight savings just changed over the weekend for London, so there has been some sort of data change. I have changed this back to normal and started a new file; 2016-03-31: See log; 2016-04-03: Makeup flow - 416.2; 2016-04-10: Zero hepa check OK
3772: N10 CNC |  |  |  | RH
Date (PC) |  | ? |  |  |  | 2016-03-16 | 2016-03-17 | 2016-03-19 | 2016-03-20 | 2016-03-20 | 2016-03-22 | 2016-03-23 | 2016-03-24 | 2016-03-25 | 2016-03-26 | 2016-03-27 | 2016-03-28 | 2016-03-29 | 2016-03-30 | 2016-03-31 | 2016-04-01 | 2016-04-02 | 2016-04-03 | 2016-04-04 | 2016-04-05 | 2016-04-06 | 2016-04-07 | 2016-04-08 | 2016-04-09 | 2016-04-10 | 2016-04-11
Time (CPC PC) | See tip above. | 22:24 |  |  |  | 23:29 | 22:35 | 12:36AM | 2:55AM | 22:21 | 1:55 | 0:47 | 2:35 | 0:53AM | 0:37 | 2:37AM | 3:42 | 3:06 | 2:48 | 6:34 | 1:39 | 1:00 | 3:14 | 2:33 | 3:30 | 0:45AM | 2:55 | 4:51 | 3:41 | 5:03 | 3:15
Check data backed up on ship servers | bar: 2016-04-03 to 2016-04-11
Sample # |  | 3 |  |  |  | 3 | 6 | 9 | 13 | 15 | 19 | 22 | 25 | 28 | 31 | 34 | 37 | 1 | 3 | 7 | 9 | 12 | 16 | 18 | 22 | 24 | 27 | 31 | 34 | 37 | 40
Concentration |  | 130 |  |  |  | 170000 | 340 | High (currently sampling pollution) | 80 | 130 | 270 | 31200 | 490 | 490 | 400 | 300 | 200 | 300000 | 470 | 120 | 400 | 350 | 110 | 400 | Pollution currently | 200 | 135 | 450 | 270 | 270 | 45
Sample graph OK? (e.g. not measuring zero or just noise) |  | Y |  |  |  | Sampling pollution |  | Ok, just pollution | Y | Y | OK | Sampling pollution | Ok, Sampling pollution occasionally | OK | Y | OK | Scale of graph off | OK, pollution | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK
CPC instrument status light yellow? (Y/N) | Only "Check flow" should be showing since the internal pump is deactivated | N, nozzle pressure |  | Y | bar: 2016-03-16 to 2016-04-11
CPC butanol Supply bottle (OK or refilled)? | Ensure there is enough butanol to last until the next check. The consumption rate will need to be determined on site. | N |  | OK |  | Ok | OK | OK | Y | OK | OK | OK | OK | N | OK | OK | OK | OK, pollution | OK | OK, refilled yesterday | OK | OK | OK | OK | OK | OK | OK | OK | OK | Refilled | OK
CPC butanol drain emptied? (Y/N) | Waste butanol cannot go down the sink. Please use an old Winchester bottle for waste and return to mainland. | N |  | N | [13 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | bar: Ok to OK
CPC status menu: Ambient pressure (should be 100 kPa) |  | 101.5 |  | 101.5 |  | 99.2 | 99.1 | 100.9 | 101.5 | 101.5 | 101 | 100.5 | 101.2 | 100.4 | 100.6 | 100.8 | 100.6 | 101 | 100.5 | 99.4 | 97.9 | 98.3 | 99.9 | 99.3 | 97.5 | 99.1 | 98.8 | 99 | 97.7 | 97.3 | 99.6
CPC Status menu: Orifice pressure (should be 10-90 kPa) |  | 79.4 |  | 77.2 |  | 75.3 | 74.9 | 76.3 | 76.9 | 77 | 76.7 | 76 | 76.3 | 76 | 76 | 75.9 | 75.7 | 76.1 | 75.8 | 74.7 | 73.7 | 73.9 | 75.3 | 74.8 | 73.4 | 74.5 | 74.3 | 74.3 | 73.4 | 72.8 | 74.7
CPC Status menu: Nozzle pressure (should be 1-6 kPa) |  | 3.1 |  | 4.12 |  | 4.044 | 4.187 | 4.384 | 4.453 | 4.493 | 4.5 | 4.469 | 4.553 | 4.484 | 4.53 | 4.591 | 4.616 | 4.64 | 4.647 | 4,57 | 4.46 | 4.53 | 4.65 | 4.57 | 4.447 | 4.602 | 4.596 | 4.667 | 4.529 | 4.591 | 4.729
Inlet flow rate (LPM or CCM) | Using gilibrator (every few days) |  |  |  |  | Ok: 1514 (23:32); OK: 1500 (0:35); OK: 1535; OK: 1546 |  |  | 1561 (at 6:51 |  |  | 1548 |  |  |  | 1501 |  |  | 1542
Zero test concentration (should be zero) | Every few days. Connect the HEPA filter to the inlet of the instrument and check the concentration falls to zero. If not, there is a leak in the system. Consult the manual and inform Ruhi. |  |  | OK |  | OK (23:32) | Ok (22:35) | OK (0:44) | OK (2:57) | OK (22:22) | OK | OK (0:48) | OK (2:38) | OK (0:55) | OK | OK (2:39) | OK (3:43) |  | OK (2:51) |  |  |  | OK |  |  |  | OK |  |  | OK
program restarted? | only do if something goes wrong |  |  | Yes, final setup performed | [13 empty] | No | No | No | No | No | No | No | No | No | No | No | No | No | No | bar: Ok to OK
New file name (if program restarted)
Comments | [5 empty] | Ok: Sampling pollution from ship exhaust; OK: sampling ship exhaust; OK: Sampling pollution; OK: Sampling pollution occasionally; OK: Can't figure out how to fix scale of computer graph | Clock time changed
3776: N3 CNC
Date (PC) |  | ? |  | ? |  | 2016-03-16 | 2016-03-17 | 2016-03-19 | 2016-03-20 | 2016-03-20 | 2016-03-22 | 2016-03-23 | 2016-03-24 | 2016-03-25 | 26/03/216 | 2016-03-27 | 2016-03-28 | 2016-03-29 | 2016-03-30 | 2016-03-31 | 2016-04-01 | 2016-04-02 | 2016-04-03 | 2016-04-04 | 2016-04-05 | 2016-04-06 | 2016-04-07 | 2016-04-08 | 2016-04-09 | 2016-04-10 | 2016-04-11
Time (CPC PC) | See tip above. | 22:26 |  | 0:20 |  | 23:25 | 22:36 | 12:44AM | 2:58AM | 22:23 | 1:56 | 0:48 | 2:38AM | 0:55AM | 0:36 | 2:39 | 3:44 | 3:11 | 2:51 | 6:36 | 1:40 | 1:01 | 3:13 | 2:34 | 3:32 | 0:46AM | 2:53 | 4:52 | 3:42 | 5:04 | 3:16
Check data backed up on ship servers | bar: 2016-04-03 to 2016-04-11
Sample # |  | 3 |  | 1 |  | 3 | 6 | 9 | 12 | 15 | 19 | 22 | 25 | 28 | 31 | 34 | 37 | 1 | 3 | 7 | 9 | 12 | 16 | 18 | 22 | 24 | 27 | 31 | 34 | 37 | 40
Concentration |  | 130 |  | 350 |  | 400000 | 380 | High (currently sampling pollution) | 90 | 136 | 270 | 33500 | 570 | 460 | 400 | 280 | 200 | 300000 | 450 | 120 | 400 | 420 | 120 | 400 | Pollution | 160 | 100 | 470 | 260 | 400 | 60
Sample graph OK? (e.g. not measuring zero or just noise) |  | OK |  | OK |  | Sampling pollution |  | Ok, just pollution | OK | OK | OK | Sampling pollution | OK, Sampling pollution occasionally | OK | OK | OK | Scale of graph off | Measuring pollution | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK
CPC instrument status bar green? (Y/N) | If not green, use the instrument display to determine the cause by entering the "Status" menu. When values are outside their recommended ranges, they will appear with "!". Record the comments here, try to fix where possible, and inform Ruhi. | Y |  | Y |  | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | bar: 2016-03-27 to 2016-04-11
CPC butanol Supply bottle (OK or refilled)? | Ensure there is enough butanol to last until the next check. The consumption rate will need to be determined on site. | OK |  | OK |  | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | Refilled | OK
CPC butanol drain emptied? (Y/N) | Waste butanol cannot go down the sink. Please use an old Winchester bottle for waste and return to mainland. | N |  | N |  | N | N | N | N | N | No | N | N | N | No | bar: 2016-03-27 to 2016-04-11
3776 CPC pressure A (~100) |  | 100.8 |  | 100.8 |  | 99.3 | 99.1 | 100.8 | 100.8 | 100.8 | 100.8 | 100.6 | 100.8 | 100.6 | 100.7 | 100.8 | 100.8 | 100.8 | 100.6 | 99.5 | 98 | 98.5 | 100.2 | 99.4 | 97.6 | 99.1 | 98.9 | 99.1 | 97.7 | 97.4 | 99.6
3776 CPC pressure O (~70) |  | 58.8 |  | 58.7 |  | 57.4 | 57.1 | 58.2 | 58.4 | 58.4 | 58.7 | 58.3 | 58.5 | 57.9 | 58 | 58 | 58.3 | 58.4 | 58 | 57.5 | 56.7 | 56.7 | 57.5 | 56.8 | 56.6 | 57.2 | 57 | 57 | 56.1 | 56.2 | 56.4
3776 CPC pressure N (~3) |  | 2.645 |  | 2.67 |  | 2.616 | 2.633 | 2.667 | 2.68 | 2.684 | 2.671 | 2.681 | 2.664 | 2.649 | 2644 | 2.644 | 2.64 | 2.65 | 2.638 | 2.6 | 2.58 | 2.59 | 2.618 | 2.607 | 2.564 | 2.602 | 2.593 | 2.593 | 2.558 | 2.547 | 2.593
Inlet flow rate | Every few days. Do at the same time as zero test (i.e. connect HEPA filter to gilibrator) |  |  |  |  | 1498 (23:34) |  |  |  | 1500 (0:35) | 1507 |  |  |  | 1508 | 2016-03-31: 1516 (at 06:52); 2016-04-03: 1509; 2016-04-07: 1547; 2016-04-10: 1501
Zero test concentration (should be zero) | Connect the HEPA filter to the inlet of the instrument and check the concentration falls to zero. If not, there is a leak in the system. Consult the manual and inform Ruhi. |  |  | OK |  | OK (23:28) | OK (22:37) | OK (0:46) | OK (2:57) | OK (22:25) | OK | OK (12:49) | OK (2:38) | OK (0:55) | OK | 2016-03-27: OK (2:39); 2016-03-28: OK (3:43); 2016-03-30: OK (2:51); 2016-03-31: OK; 2016-04-03: OK; 2016-04-07: OK; 2016-04-10: OK
program restarted? | only do if something goes wrong |  |  | Yes, final setup running |  | No | No | No | No | No | No | No | No | No | No | bar: 2016-03-27 to 2016-03-28; 2016-03-30 to 2016-03-31; 2016-04-03 to 2016-04-11
New file name (if program restarted)
Comments | [5 empty] | Sampling pollution from ship exhaust |  | Sampling pollution from ship exhaust |  |  |  | Sampling pollution | Sampling pollution occasionally |  |  | 2016-03-28: Refill butanol soon?
GRIMM nanoSMPS
Date (PC) |  | ? | ? | ? |  | 2016-03-16 | 2016-03-17 | 2016-03-19 | 2016-03-20 | 2016-03-19 | 2016-03-22 | 2016-03-23 | 2016-03-24 | 2016-03-15 | 2016-03-26 | 2016-03-27 | 2016-03-28 | 2016-03-29 | 2016-03-30 | 2016-03-31 | 2016-04-01 | 2016-04-02 | 2016-04-03 | 2016-04-04 | 2016-04-05 | 2016-04-06 | 2016-04-07 | 2016-04-08 | 2016-04-09 | 2016-04-10 | 2016-04-11
Time (PC) |  | 22:24 | 3:12 | 10:54 |  | 23:21 | 22:38 | 12:47AM | 3:01AM | 22:11 | 2:01 | 0:50 | 2:40AM | 0:56AM | 0:42 | 1:42AM | 2:37AM | 3:12 | 2:53AM | 6:38 | 1:41 | 1:02 | 3:17 | 2:26 | 3:34 | 0:47AM | 2:52 | 4:53 | 3:43 | 4:45 | 3:17
any errors? |  | N | N | N |  | N | N | N | N | N | N | N | 00256 (23/03/16 05:39:39) | N | N | N | N | N | N | N | N | N | N | N | N | 8 | N | N | N | N | N
total CN conc. (after scan is finished) |  | 94 | 185 | 215 |  | 173801 | 152 | 71855 | 49 | 53 |  | 11103 | 2314 | 257 | 200 | 147 | 89 |  | 194 | 60 | 250 | 80 | 54 | 150 | 400 | 145 | 59 | 128 | 120 | 74 | 35
status LED's OK? |  | Y | Y | Y |  | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | bar: 2016-03-27 to 2016-04-11
butanol Supply bottle OK? |  | OK | OK | OK |  | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | K | OK | Y | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK
Inlet flow rate | Every few days. Turn tap off to manifold before unplugging from inlet, and MAKE SURE TO TURN THE TAP BACK ON WHEN RECONNECTING | [13 empty] | 276.4 | 2016-04-03: 282.1; 2016-04-10: 280.8
Comments | [15 empty] | 2016-03-28: Refill butanol soon?; 2016-03-30: Refilled; 2016-04-06: Software and hardware restarted due to butanol feed error
Polar Nephelometer
Date (PC) |  | ? | ? | ? |  | 2016-03-17 | 2016-03-18 | 2016-03-19 | 2016-03-20 | 2016-03-21 | 2016-03-22 | 2016-03-23 | 2016-03-24 | 2016-03-25 | 2016-03-26 | 2016-03-27 | 2016-03-28 | 2016-03-29 | 2016-03-30 | 2016-03-31 | 2016-04-01 | 2016-04-02 | 2016-04-03 | 2016-04-04 | 2016-04-05 | 2016-04-06 | 2016-04-07 | 2016-04-08 | 2016-04-09 | 2016-04-10 | 2016-04-11
Neph clock time | Should be on Hobart Time | 9:25 | 14:06 | 9:45 |  | 10:14 | 9:42 | 11:50AM | 14:02 | 09:26AM | 13:04 | 11:56 | 13:41 | 11:58AM | 11:50 | 12:43PM | 13:45 | 14:13 | 13:54 | 17:40 | 12:47pm | 12:03 | 14:19 | 13:24 | 14:03 | 11:58 | 13:47 | 15:53 | 14:44 | 15:41 | 14:20
station time | [5 empty] | 10:14AM | 9:42 | 11:50AM | 14:02 | 09:27AM | 12:04 | 11:56 | 13:42 | 11:59AM | 10:51 | 12:42PM | 13:46 | 13:14 | 13:55 | 16:40 | 11:47 | 12:03 | 13:19 | 12:24 | 13:03 | 11:00 | 12:49 | 14:55 | 13:43 | 14:43 | 13:20
Scatt (3 WL's) | 635/525/450 nm at 0 degrees (take a photo and copy from there) | 1.3/0.7/1.5 | 6.4/5.3/4.5 | 6.4/5.3/2.5 | [6 empty] | 5.7/5.3/6.4 |  |  |  | 35.6/35.1/28 | 17.9/15.7/15.8 | 13.1/11.1/11.3 | 44/56/68 | 23.7/21.4/20.9 | 18.2/16.5/15.4 | 48.6/40.9/40.7 | 7.73/6.32/6.87 | 27.6/24.3/23.6 | 14.748/12.860/12.557 | 66.711/54.791/55.953 | 12.9/11.4/11.7 | 32.3/29.928.4 | 27.2/22.7/23.9 | 23.785/21.278/19.865 | 13.9/11.7/12.0 | 4.429/2.882/3.871
last Zero check | record the Zero check value for each of the 3 wavelengths. These are in the Calibration menu, then go into 'Parameter (wavelength) 1...3' and record the 'Last Zero' value. |  |  |  |  | 2016-03-31: 2.32/1.7828/2.7047; 2016-04-01: 23.41/20.49/19.308; 2016-04-02: 7.6155/7.1543/8.0964; 2016-04-03: 2.5152/1.9557/2/3923; 2016-04-04: 4.0132/3.4313/4.5887; 2016-04-05: 28.497/24.026/23.530; 2016-04-06: 7.0809/5.9595/6.0139; 2016-04-07: 3.0415/1.8766/3.2593; 2016-04-08: 24.969/22.475/22.331; 2016-04-09: 33.623/31.705/28.936; 2016-04-10: 9.3065/6.8358/7.3388; 2016-04-11: 5.2665/4.9384/5.9796
Cell RH (<40%) |  | 39 | 37 | 40 |  | 39 | 28 | 26 | 39 | 38 | 39 | 38 | 30 | 35 | 34 | 34 | 31 | 36 | 19 | 37 | 40 | 27 | 34 | 35 | 40 | 20 | 29 | 25 | 31 | 24 | 22
Cell Temp |  | 29.7 | 25.2 | 22.7 |  | 26.2 | 21.7 | 24.7 | 22.8 | 22.4 | 24 | 26.7 | 22.6 | 23 | 22.8 | 23.1 | 23.4 | 22.7 | 23.5 | 25.1 | 23.9 | 24.4 | 23.2 | 24.3 | 23.4 | 25.3 | 25.2 | 22.6 | 24 | 22.6 | 24.5
exhaust fan/pump ON? (feel) |  | Y | Y | Y | bar: 2016-03-17 to 2016-04-01; 2016-04-03 to 2016-04-11
Comments
Thermo MAAP BC
Date (PC) | [19 empty] | 2016-03-31 | 2016-04-01 | 2016-04-02 | 2016-04-03 | 2016-04-04 | 2016-04-05 | 2016-04-06 | 2016-04-07 | 2016-04-08 | 2016-04-09 | 2016-04-10 | 2016-04-11
Instrument clock | From the Main Screen, press <disp> and <next> until you see the Clock screen. Exit by pressing <next> until the Main Screen. | [18 empty] | 16:50 | 11:46 | 11:09 | 13:26 | 12:40 | 13:05 | 11:05 | 12:50 | 15:00 | 13:53 | 14:34 | 13:21
Station time | [19 empty] | 16:46 | 11:42 | 11:06 | 13:22 | 12:37 | 13:01 | 11:03 | 12:48 | 14:58 | 13:49 | 14:32 | 13:18
BC conc | ug/m3 | 0.38 | 0 | -0.04 |  | 8.57 | -0.04 | 7.63 | 0 | -0.06 | -0.01 | -0.02 | 5.01 | 0.03 | 0 | 0.19 | 0.05 | 11.83 | 0.09 | -0.05 | 0 | 0 | 0 | 0 | 0 | -0.01 | 0.05 | 0 | 0 | 4.48 | 0
status code |  | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
error code |  | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
displayed flow (1000L/hr) |  | 998 | 1000 | 1000 |  | 1002 | 1000 | 1000 | 1001 | 1000 | 1000 | 1000 | 1002 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 999 | 999 | 1000 | 1000 | 1000 | 1000
T1…T4 temps | From the Main Screen (Home screen) press <next> until the 'Service Menu' is displayed, then <yes> and make sure 'Online' is displayed. If you <set> this and change it to 'offline/enable keypad' then data output will cease. Once in the Service menu, press <NEXT> to scroll through the screens until the Temps are displayed. Unfortunately this is not talked about in the Ops Manual, but it should be the next screen after the 'ANALOG INPUT IN LSB' screen. Once they are recorded, press <back> or <next> until you get back to the Main Screen. | ? | ? | ? |  | ? | ? | ? | ? | ? |  | ? | ? | ? | [5 empty] | 21/28 | 20/27.5 | 21/27.6 | 19/28 | 20/28 | 18/28 | 23/28 | 22/27.6 | 19/27.2 | 21/27.4 | 18/27.6 | 21/28
comments | [19 empty] | 2016-04-01: Checked instrument clock offset with technies and IT and the data is timestamped by the system, rather than the instrument, so this isn't an issue.
Nafion
Sheath flow rate | Rotameter mounted on starboard side of aerosol lab, on the right of the rack. Flowrate should be 5 LPM. Change if necessary. | Slightly low, changed to 5LPM | Steady at 5 LPM | Steady |  | Steady | Steady 4.5-5.5 | Steady 4.5-5.5 | Steady 5 | Steady 5 | Steady 5 | Steady at 4.5 | Steady at 4.5 | Steady at 4.5 | Adjusted from 4.5 to 5 | Steady 4-4.5 | Steady 4.5-5 | Slightly low, adjusted back up to 5 | 5 | 5 steady | 5 steady | 5 steady | 5 steady | 5 steady | 5 steady | 5 steady | 5 steady | 5 steady | 5 steady | 5 steady | 5 steady
WEEKLY
RH downstream of nafion | Use external RH probe and insert into the flow downstream of nafion |  |  | 14% |  | 56.3 | 44.5 | 35 | 53.4 | 53.3 | 53.8 | 64.2 | 41.3 | 51.1 | 47.1 | 43.4 | 41.1 | 47 | 27.1 | 43.8 | 51.3 | 32.4 | 44 | 45.9 | 54.3 | 21.4 | 35.1 | 30.8 | 39.5 | 29.1 | 23.3
RH upstream in manifold | As per underway dataset |  |  | 75.4 |  | 96 | 66 |  | 84 | 86 | 69.5 | 86 | 53 | 64 |  |  | 73 | 66 | 56 | 83.2 |  | 68.3 |  |  | 92 | 66 | 74 | 56.9 |  | 59 | 61
Check dessicant colour | If wet, replace and dry wet in 40 or 60oC oven |  |  |  |  | Ok | OK | OK | OK | Needs changing | OK | OK | OK | OK | OK, change tomorrow | Needs changing |  | OK | OK | OK | OK | OK | OK | OK |  | Needs changing | OK | OK | OK | Needs changing | OK
Comments |  |  |  | Changed dessicant | [5 empty] | Dessicant changed (09:30AM station time) | [5 empty] | Dessicant changed (14:00 station time)
Environment
Lab temperature |  |  |  | 23.5
Lab RH |  |  |  | 43%
general check Intake lines | Every few days |  |  | OK | bar: 2016-04-03
water in Manifold? | Use torch to look through window on manifold |  |  | no |  | No | No | No | No | No |  | No, but window is foggy | No, but window is foggy | No, but window is foggy | OK | OK | OK |  | OK |  |  | No | No |  | Yes, a fair bit. And I think it has leaked down from the roof too, because there is water droplets and salt on lots of the lines around, as well as the APS | No | No | No | Damp, but empty | No | Yes, only minimal though
Manifold (x2) Pressure checks | a few times on voyage, in each lab
Comments
MOUDI
Check pressure gauge
Weekly | bar: 2016-04-05
Change filters