## CAPRICORN_Aerosols
Date |  | 2016-03-13 | 2016-03-14 | 2016-03-15 | 2016-03-16 | 2016-03-17 | 2016-03-18 | 2016-03-19 | 2016-03-20 | 2016-03-21 | 2016-03-22 | 2016-03-23 | 2016-03-24 | 2016-03-25 | 2016-03-26 | 2016-03-27 | 2016-03-28 | 2016-03-29 | 2016-03-30 | 2016-03-31 | 2016-04-01 | 2016-04-02 | 2016-04-03 | 2016-04-04 | 2016-04-05 | 2016-04-06 | 2016-04-07 | 2016-04-08 | 2016-04-09 | 2016-04-10 | 2016-04-11 | 2016-04-12 | 2016-04-13 | 2016-04-14 | 2016-04-15
Operator |  | 2016-03-13: RH; 2016-03-14: KL; 2016-03-15: KL; 2016-03-17: KL; 2016-03-18: KL; 2016-03-19: RH/KL; 2016-03-20: KL; 2016-03-21: KL
APS
Time (station) | Our computer time (GMT+10) | 2016-03-17: 10:09AM; 2016-03-18: 09:23AM; 2016-03-19: 11:24AM; 2016-03-20: 13:49; 2016-03-21: 09:16AM
Date (PC) |  | 2016-03-13: ?; 2016-03-14: ?; 2016-03-15: ?; 2016-03-17: 2016-03-16; 2016-03-18: 2016-03-17; 2016-03-19: 2016-03-19; 2016-03-20: 2016-03-20; 2016-03-21: 2016-03-20
Time (PC) | If this is different to station time, please synchronise clocks - should be set to sync with internet time. | 2016-03-13: 22:05; 2016-03-14: 14:54; 2016-03-15: 9:51; 2016-03-17: 11:08PM; 2016-03-18: 22:24; 2016-03-19: 12:23AM; 2016-03-20: 2:50AM; 2016-03-21: 22:15
Check data backed up on ship servers
Sample distribution OK? |  | 2016-03-13: Y; 2016-03-14: Y; 2016-03-15: Y; 2016-03-17: Y; 2016-03-18: Y; 2016-03-19: Y; 2016-03-20: Y; 2016-03-21: A bit noisy
Sheath flow (LPM) | In display menu | 2016-03-13: 4; 2016-03-14: 4; 2016-03-15: 4; 2016-03-17: 4; 2016-03-18: 4; 2016-03-19: 4; 2016-03-20: 4; 2016-03-21: 4
Inlet flow (LPM) | In display menu | bar: 2016-03-13 to 2016-03-15; 2016-03-17 to 2016-03-21
Total flow (LPM) | Should be the sum of sheath and inlet | 2016-03-13: 5; 2016-03-14: 5; 2016-03-15: 5; 2016-03-17: 5; 2016-03-18: 5; 2016-03-19: 5; 2016-03-20: 5; 2016-03-21: 5
Laser power (%) | In display menu | 2016-03-13: 75; 2016-03-14: 75; 2016-03-15: 75; 2016-03-17: 75; 2016-03-18: 75; 2016-03-19: 75; 2016-03-20: 75; 2016-03-21: 75
Laser current (mA) |  | 2016-03-13: 52.4; 2016-03-14: 53.2; 2016-03-15: 52; 2016-03-17: 52.3; 2016-03-18: 51.6; 2016-03-19: 53.2; 2016-03-20: 52.6; 2016-03-21: 52.5
Comments |  | 2016-03-13: Very low concentrations, misty/foggy day with rain this morning; 2016-03-19: Low concentrations
CCNC
Time (station) | Our computer time (GMT+10)
Date (PC) | bar: 2016-03-13 to 2016-03-15
Time (CCNC PC) | If this is different to station time, please synchronise clocks - should be set to sync with internet time. | 2016-03-13: 22:11; 2016-03-14: 14:58; 2016-03-15: 10:49
Check data backed up on ship servers
Status box green? (Y/N) | If "NO", and it is not an obvious fix, please email Ruhi.Humphries@csiro.au | bar: 2016-03-13 to 2016-03-15
Size distribution Gaussian? If not, please describe |  | 2016-03-13: Yes, but not a great gaussian, likely due to the low concentrations; 2016-03-14: Yes, but not a great gaussian, likely due to the low concentrations; 2016-03-15: Y
SS tab: current SS |  | 2016-03-13: 0.5; 2016-03-14: 0.5; 2016-03-15: 0.5
Flow tab: Flow ratio (should be between 10.0 ± 0.1) | If not within the range, it may be a small adjustment in the sheath valve. If not, a flow calibration may be necessary. Please consult manual and inform Ruhi. | bar: 2016-03-13 to 2016-03-15
OPC tab: Concentration |  | 2016-03-13: 20; 2016-03-14: 44; 2016-03-15: 173
CCNC water reservoir (mL) | The CCNC column must NOT dry out. Please ensure there is enough milli-Q water present until next check. | 2016-03-13: Full; 2016-03-14: 2/3 Full; 2016-03-15: Full
Water refilled and drain emptied? (Y/N) | Once refilled, ensure all fittings are secured tightly as this is the most common location for leaks to be introduced. Ensure blue cap and white fittings in each cap are tight. Only use Milli-Q water, and ensure no dirt enters the water during refilling. Drain bottles can be emptied down the sink. | bar: 2016-03-13 to 2016-03-15
OPC water trap (comments) | Check OPC water trap. If liquid is present, empty the bottle. | bar: 2016-03-13 to 2016-03-15
Check bottom of case for water leakage | bar: 2016-03-13 to 2016-03-15
Zero test concentration (should be <5/sec) | Check after changing the water | bar: 2016-03-15
Comments
WEEKLY | bar: 2016-03-15; 2016-03-22
Measure inlet flow using external flowmeter | Inlet flow should be 500 ccm (0.5 LPM) | 2016-03-15: Tested at 11:13pm UTC at 0.494 LPM
TSI long SMPS
Time (station) | Our computer time (GMT+10) | 2016-03-17: 10:35; 2016-03-18: 09:27AM; 2016-03-19: 11:27AM; 2016-03-20: 13:52; 2016-03-21: 09:17AM
Date (PC) |  | 2016-03-15: ?; 2016-03-17: 2016-03-16; 2016-03-18: 2016-03-17; 2016-03-19: 2016-03-19; 2016-03-20: 2016-03-20; 2016-03-21: 2016-03-20
Time (PC) | If this is different to station time, please synchronise clocks - should be set to sync with internet time. | 2016-03-15: 0:24; 2016-03-17: 23:35; 2016-03-18: 22:26; 2016-03-19: 12:27AM; 2016-03-20: 2:52AM; 2016-03-21: 22:17
Check data backed up on ship servers
Sample # |  | 2016-03-18: 207; 2016-03-19: 516; 2016-03-20: 833; 2016-03-21: 1066
Sample distribution OK? | Please describe if not normal | 2016-03-17: OK; 2016-03-18: OK; 2016-03-19: OK; 2016-03-20: OK; 2016-03-21: Quite noisy
Zero counts on CPC between scans? |  | 2016-03-17: 0.06; 2016-03-19: 0; 2016-03-20: 0; 2016-03-21: 0
SMPS status light green | bar: 2016-03-17 to 2016-03-21
CPC butanol reservoir level (OK or filled)? | bar: 2016-03-17 to 2016-03-21
3776 CPC pressure A (~100) | On CPC screen | 2016-03-15: 101.5; 2016-03-17: 100.4; 2016-03-18: 100.3; 2016-03-19: 101.5; 2016-03-20: 101.5; 2016-03-21: 101.5
3776 CPC pressure O (~70) | On CPC screen | 2016-03-15: 79.6; 2016-03-17: 77.5; 2016-03-18: 77.4; 2016-03-19: 79; 2016-03-20: 79.5; 2016-03-21: 79.7
3776 CPC pressure N (~3) | On CPC screen | 2016-03-15: 3.1; 2016-03-17: 3; 2016-03-18: 3; 2016-03-19: 3.1; 2016-03-20: 3.1; 2016-03-21: 3.1
Flows: Sheath inline (3 LPM) | Check on screen | 2016-03-19: 3; 2016-03-20: 3; 2016-03-21: 3
Flows: Monodisperse inline | Every few days (1L before, 300mL after the 18th) | 2016-03-17: 990.4 (23:42); 2016-03-21: ?? (23:11)
Flows: polydisperse in | Every few days (1L before, 300mL after the 18th) | 2016-03-17: 986.8 (23:45); 2016-03-21: 439 (23:06)
Flows: Inlet flow (1LPM) | Every few days
program restarted? | only do if something goes wrong | bar: 2016-03-17 to 2016-03-21
New file name (if program restarted)
Comments
3772: N10 CNC | bar: 2016-03-15
Date (PC) |  | 2016-03-13: ?; 2016-03-17: 2016-03-16; 2016-03-18: 2016-03-17; 2016-03-19: 2016-03-19; 2016-03-20: 2016-03-20; 2016-03-21: 2016-03-20
Time (CPC PC) | See tip above. | 2016-03-13: 22:24; 2016-03-17: 23:29; 2016-03-18: 22:35; 2016-03-19: 12:36AM; 2016-03-20: 2:55AM; 2016-03-21: 22:21
Check data backed up on ship servers
Sample # |  | 2016-03-13: 3; 2016-03-17: 3; 2016-03-18: 6; 2016-03-19: 9; 2016-03-20: 13; 2016-03-21: 15
Concentration |  | 2016-03-13: 130; 2016-03-17: 170000; 2016-03-18: 340; 2016-03-19: High (currently sampling pollution); 2016-03-20: 80; 2016-03-21: 130
Sample graph OK? (e.g. not measuring zero or just noise) |  | 2016-03-13: Y; 2016-03-17: Sampling pollution; 2016-03-19: Ok, just pollution; 2016-03-20: Y; 2016-03-21: Y
CPC instrument status light yellow? (Y/N) | Only "Check flow" should be showing since the internal pump is deactivated | 2016-03-13: N, nozzle pressure; 2016-03-15: Y; 2016-03-17: Y; 2016-03-18: Y; 2016-03-19: Y; 2016-03-20: Y; 2016-03-21: Y
CPC butanol Supply bottle (OK or refilled)? | Ensure there is enough butanol to last until the next check. The consumption rate will need to be determined on site. | 2016-03-13: N; 2016-03-15: OK; 2016-03-17: Ok; 2016-03-18: OK; 2016-03-19: OK; 2016-03-20: Y; 2016-03-21: OK
CPC butanol drain emptied? (Y/N) | Waste butanol cannot go down the sink. Please use an old Winchester bottle for waste and return to mainland. | bar: 2016-03-13; 2016-03-15; 2016-03-17 to 2016-03-21
CPC Status menu: Orifice pressure (should be 10-90 kPa) |  | 2016-03-13: 79.4; 2016-03-15: 77.2; 2016-03-17: 75.3; 2016-03-18: 74.9; 2016-03-19: 76.3; 2016-03-20: 76.9; 2016-03-21: 77
CPC Status menu: Nozzle pressure (should be 1-6 kPa) |  | 2016-03-13: 3.1; 2016-03-15: 4.12; 2016-03-17: 4.044; 2016-03-18: 4.187; 2016-03-19: 4.384; 2016-03-20: 4.453; 2016-03-21: 4.493
CPC status menu: Ambient pressure (should be 100 kPa) |  | 2016-03-13: 101.5; 2016-03-15: 101.5; 2016-03-17: 99.2; 2016-03-18: 99.1; 2016-03-19: 100.9; 2016-03-20: 101.5; 2016-03-21: 101.5
Inlet flow rate (LPM or CCM) | Using gilibrator (every few days) | 2016-03-17: 1514 (23:32)
Zero test concentration (should be zero) | Connect the HEPA filter to the inlet of the instrument and check the concentration falls to zero. If not, there is a leak in the system. Consult the manual and inform Ruhi. | 2016-03-15: OK; 2016-03-17: OK (23:32); 2016-03-18: Ok (22:35); 2016-03-19: OK (0:44); 2016-03-20: OK (2:57); 2016-03-21: OK (22:22)
program restarted? | only do if something goes wrong | 2016-03-15: Yes, final setup performed; 2016-03-17: No; 2016-03-18: No; 2016-03-19: No; 2016-03-20: No; 2016-03-21: No
New file name (if program restarted)
Comments |  | 2016-03-17: Sampling pollution from ship exhaust; 2016-03-19: sampling ship exhaust
3776: N3 CNC
Date (PC) |  | 2016-03-13: ?; 2016-03-15: ?; 2016-03-17: 2016-03-16; 2016-03-18: 2016-03-17; 2016-03-19: 2016-03-19; 2016-03-20: 2016-03-20; 2016-03-21: 2016-03-20
Time (CPC PC) | See tip above. | 2016-03-13: 22:26; 2016-03-15: 0:20; 2016-03-17: 23:25; 2016-03-18: 22:36; 2016-03-19: 12:44AM; 2016-03-20: 2:58AM; 2016-03-21: 22:23
Check data backed up on ship servers
Sample # |  | 2016-03-13: 3; 2016-03-15: 1; 2016-03-17: 3; 2016-03-18: 6; 2016-03-19: 9; 2016-03-20: 12; 2016-03-21: 15
Concentration |  | 2016-03-13: 130; 2016-03-15: 350; 2016-03-17: 400000; 2016-03-18: 380; 2016-03-19: High (currently sampling pollution); 2016-03-20: 90; 2016-03-21: 136
Sample graph OK? (e.g. not measuring zero or just noise) |  | 2016-03-13: OK; 2016-03-15: OK; 2016-03-17: Sampling pollution; 2016-03-19: Ok, just pollution; 2016-03-20: OK; 2016-03-21: OK
CPC instrument status bar green? (Y/N) | If not green, use the instrument display to determine the cause by entering the "Status" menu. When values are outside their recommended ranges, they will appear with "!". Record the comments here, try to fix where possible, and inform Ruhi. | bar: 2016-03-13; 2016-03-15; 2016-03-17 to 2016-03-21
CPC butanol Supply bottlel (OK or refilled)? (Y/N) | Ensure there is enough butanol to last until the next check. The consumption rate will need to be determined on site. | bar: 2016-03-13; 2016-03-15; 2016-03-17 to 2016-03-21
CPC butanol Supply refilled and drain bottle emptied? (Y/N) | Ensure that refilling is performed in a well-ventilated area. | bar: 2016-03-13; 2016-03-15; 2016-03-17 to 2016-03-21
CPC butanol drain emptied? (Y/N) | Waste butanol cannot go down the sink. Please use an old Winchester bottle for waste and return to mainland. | bar: 2016-03-13; 2016-03-15; 2016-03-17 to 2016-03-21
3776 CPC pressure A (~100) |  | 2016-03-13: 100.8; 2016-03-15: 100.8; 2016-03-17: 99.3; 2016-03-18: 99.1; 2016-03-19: 100.8; 2016-03-20: 100.8; 2016-03-21: 100.8
3776 CPC pressure O (~70) |  | 2016-03-13: 58.8; 2016-03-15: 58.7; 2016-03-17: 57.4; 2016-03-18: 57.1; 2016-03-19: 58.2; 2016-03-20: 58.4; 2016-03-21: 58.4
3776 CPC pressure N (~3) |  | 2016-03-13: 2.645; 2016-03-15: 2.67; 2016-03-17: 2.616; 2016-03-18: 2.633; 2016-03-19: 2.667; 2016-03-20: 2.68; 2016-03-21: 2.684
Inlet flow rate | Every few days | 2016-03-17: 1498 (23:34)
Zero test concentration (should be zero) | Connect the HEPA filter to the inlet of the instrument and check the concentration falls to zero. If not, there is a leak in the system. Consult the manual and inform Ruhi. | 2016-03-15: OK; 2016-03-17: OK (23:28); 2016-03-18: OK (22:37); 2016-03-19: OK (0:46); 2016-03-20: OK (2:57); 2016-03-21: OK (22:25)
program restarted? | only do if something goes wrong | 2016-03-15: Yes, final setup running; 2016-03-17: No; 2016-03-18: No; 2016-03-19: No; 2016-03-20: No; 2016-03-21: No
New file name (if program restarted)
Comments |  | 2016-03-17: Sampling pollution from ship exhaust; 2016-03-19: Sampling pollution from ship exhaust
GRIMM nanoSMPS
Date (PC) |  | 2016-03-13: ?; 2016-03-14: ?; 2016-03-15: ?; 2016-03-17: 2016-03-16; 2016-03-18: 2016-03-17; 2016-03-19: 2016-03-19; 2016-03-20: 2016-03-20; 2016-03-21: 2016-03-19
Time (PC) |  | 2016-03-13: 22:24; 2016-03-14: 3:12; 2016-03-15: 10:54; 2016-03-17: 23:21; 2016-03-18: 22:38; 2016-03-19: 12:47AM; 2016-03-20: 3:01AM; 2016-03-21: 22:11
any errors? | bar: 2016-03-13 to 2016-03-15; 2016-03-17 to 2016-03-21
total CN conc. (after scan is finished) |  | 2016-03-13: 94; 2016-03-14: 185; 2016-03-15: 215; 2016-03-17: 173801; 2016-03-18: 152; 2016-03-19: 71855; 2016-03-20: 49; 2016-03-21: 53
status LED's OK? | bar: 2016-03-13 to 2016-03-15; 2016-03-17 to 2016-03-21
butanol Supply bottle OK? | bar: 2016-03-13 to 2016-03-15; 2016-03-17 to 2016-03-21
Comments
Polar Nephelometer
Date (PC) |  | 2016-03-13: ?; 2016-03-14: ?; 2016-03-15: ?; 2016-03-17: 2016-03-17; 2016-03-18: 2016-03-18; 2016-03-19: 2016-03-19; 2016-03-20: 2016-03-20; 2016-03-21: 2016-03-21
Neph clock time |  | 2016-03-13: 9:25; 2016-03-14: 14:06; 2016-03-15: 9:45; 2016-03-17: 10:14; 2016-03-18: 9:42; 2016-03-19: 11:50AM; 2016-03-20: 14:02; 2016-03-21: 09:26AM
station time |  | 2016-03-17: 10:14AM; 2016-03-18: 9:42; 2016-03-19: 11:50AM; 2016-03-20: 14:02; 2016-03-21: 09:27AM
Scatt (3 WL's) | 635/525/450 nm at 0 degrees | 2016-03-13: 1.3/0.7/1.5; 2016-03-14: 6.4/5.3/4.5; 2016-03-15: 6.4/5.3/2.5; 2016-03-17: 6.4/5.3/4.5; 2016-03-18: 6.4/5.3/4.5; 2016-03-19: 6.4/5.3/4.5; 2016-03-20: 6.4/5.3/4.5; 2016-03-21: 6.4/5.3/4.5
last Zero check
Cell RH (<40%) |  | 2016-03-13: 39; 2016-03-14: 37; 2016-03-15: 40; 2016-03-17: 39; 2016-03-18: 28; 2016-03-19: 26; 2016-03-20: 39; 2016-03-21: 38
Cell Temp |  | 2016-03-13: 29.7; 2016-03-14: 25.2; 2016-03-15: 22.7; 2016-03-17: 26.2; 2016-03-18: 21.7; 2016-03-19: 24.7; 2016-03-20: 22.8; 2016-03-21: 22.4
exhaust fan/pump ON? (feel) | bar: 2016-03-13 to 2016-03-15; 2016-03-17 to 2016-03-21
Comments
Thermo MAAP BC
Date (PC)
Instrument clock
BC conc | ug/m3 | 2016-03-13: 0.38; 2016-03-14: 0; 2016-03-15: -0.04; 2016-03-17: 8.57; 2016-03-18: -0.04; 2016-03-19: 7.63; 2016-03-20: 0; 2016-03-21: -0.06
status code |  | 2016-03-13: 0; 2016-03-14: 0; 2016-03-15: 0; 2016-03-17: 0; 2016-03-18: 0; 2016-03-19: 0; 2016-03-20: 0; 2016-03-21: 0
error code |  | 2016-03-13: 0; 2016-03-14: 0; 2016-03-15: 0; 2016-03-17: 0; 2016-03-18: 0; 2016-03-19: 0; 2016-03-20: 0; 2016-03-21: 0
displayed flow (1000L/hr) |  | 2016-03-13: 998; 2016-03-14: 1000; 2016-03-15: 1000; 2016-03-17: 1002; 2016-03-18: 1000; 2016-03-19: 1000; 2016-03-20: 1001; 2016-03-21: 1000
T1…T4 temps | bar: 2016-03-13 to 2016-03-15; 2016-03-17 to 2016-03-21
comments
Nafion
Sheath flow rate | Rotameter mounted on starboard side of aerosol lab, on the right of the rack. Flowrate should be 5 LPM | 2016-03-13: Slightly low, changed to 5LPM; 2016-03-14: Steady at 5 LPM; 2016-03-15: Steady; 2016-03-17: Steady; 2016-03-18: Steady 4.5-5.5; 2016-03-19: Steady 4.5-5.5; 2016-03-20: Steady 5; 2016-03-21: Steady 5
WEEKLY | bar: 2016-03-15; 2016-03-22
RH downstream of nafion | Use external RH probe and insert into the flow downstream of nafion | 2016-03-15: 14%; 2016-03-17: 56.3; 2016-03-18: 44.5; 2016-03-19: 35; 2016-03-20: 53.4; 2016-03-21: 53.3
RH upstream in manifold | As per underway dataset | 2016-03-15: 75.4; 2016-03-17: 96; 2016-03-18: 66; 2016-03-20: 84; 2016-03-21: 86
Check dessicant colour | Some blue present | 2016-03-17: Ok; 2016-03-18: OK; 2016-03-19: OK; 2016-03-20: OK; 2016-03-21: Needs changing
Comments |  | 2016-03-15: Changed dessicant; 2016-03-21: Dessicant changed (09:30AM station time)
Environment
Lab temperature |  | 2016-03-15: 23.5
Lab RH |  | 2016-03-15: 43%
general check Intake lines | Every few days | bar: 2016-03-15
water in Manifold? | bar: 2016-03-15; 2016-03-17 to 2016-03-21
Manifold (x2) Pressure checks | a few times on voyage, in each lab
Comments
MOUDI
Check pressure gauge
Weekly | bar: 2016-03-15; 2016-03-22; 2016-03-29; 2016-04-05; 2016-04-12
Change filters