L.Bariteau

April 10, 2008

Readme for PMS_10mn_GASEXIII_2008.txt and PMS_hr_GASEXIII_2008.txt

The data file PMS_10mn_GASEXIII_2008.txt contains measurements of atmospheric aerosol counts.  The instrument used is a Particle Measurement Systems (PMS) Lasair-II aerosol spectrometer.  The Lasair-II draws air through an intake and uses scatter of laser light from individual particles to determine the size.  Particles are counted in six size bins:  0.1-0.2, 0.2-0.3, 0.3-0.5, 0.5-1, 1-5, and greater than 5.0 μm diameters.  The PSD system was mounted in the University of Hawaii’s seatainer on the 02 deck with the intake on the upwind side of the container.  The system ran at 1.0 cfm (2.83e4 cm3/min) sample volume flow rate with a count deconcentrator that reduces the counts a factor of 10 (to prevent coincidence errors).

The files contain 10min and hourly aerosol averaged from 6 sec samples.  Since 6 sec is 0.1 minute, the aerosol concentration is related to the average counts obtained in 1-min by

Concentration = 10*Counts/(2.83*1e4*0.1)  (#/cm3)

The factor of 10 is the deconcentrator and the factor of 0.1 is the 6-sec sample.  The aerosol spectral concentration is dN/dr=C/dr where dr is the width of the size bin.

The columns in the file are:

1		Julian Day (2008)

2		Total counts

3		Counts in channel 0 (0.1 – 0.2 micron)

4		Counts in channel 1 (0.2 – 0.3 micron)

5		Counts in channel 2 (0.3 – 0.5 micron)

6		Counts in channel 3 (0.5 – 1 micron)

7		Counts in channel 4 (1 – 5 micron)

8		Counts in channel 5 (5 – up micron)

Figure 1.   Aerosol concentrations from Lasair-II spectrometer from JD 62 to JD 101 (March 2nd to April 10th, 2008). Upper panel: total number concentration for aerosols larger than 0.1 micron diameter. Spikes are caused by the ship’s exhaust.  Upper panel: aerosol concentrations for 0.1-0.2 (blue), 0.2-0.3 (red), 0.3-0.5 (green), 0.5-1.0 (yellow), and 1.0-5.0 (cyan).  Spikes are caused by ship’s exhaust.