Summery GOMECC 2007 (Gulf of Mexico and East Coast Carbon Cruise)

This summery will sum up observations, problems, etc during the GOMECC cruise in 2007. The cruise was going from Galveston, TX to Boston, MA from July 10th to August 4th.

Also some data analysis results will be shown here, especially an overview on the calibrations and the lag-time test will be given so that these and other important information will be available for people working on the cruise data but didn’t participate.

Instructions: For additional information on the instrumentation and instructions how to run the instrument read ‘San Diego’, ‘Directions for operating the Fast Ozone Sensor.doc’, and Diploma thesis ‘Ozone flux measurements during the Gulf of Mexico and East Coast Carbon Cruise 2007 (GOMECC)’.

- Where to find the data at NOAA

Data for the GOMECC cruise can be found at ‘R:\GOMECC_2007\KathrinsData’.

Comments on subfolders:

- ‘correction excel’ 	files that have been changed so that they can be run in MATLAB; the original files can be found in ‘R:\GOMECC_2007\RHB\flux\Raw’.

- ‘SEP08_final_for_paper’	this folder is divided into ‘48s’ and ‘52s’ which means that 
fluxes were once calculated with a lag-time of 4.8s and once with 5.2s, 
respectively. The data of the folder ‘48s’ is the data used for the paper. Smoothing correction is applied.

- ‘June08’	Correction in error of ship motion correction calculation; looking at 
		quenching effect (simulation + correct O3 signal for quenching before 
		calculating fluxes

- ‘OCT08’	Ludovic’s quenching and Webb correction applied. Unnecessary, since 
		Nafion® drier was part of instrumentation

- ‘MATLAB_general’		matlab files for corrections such as smoothing correction

- ‘Excel and Word files’	Analysis, overviews, calibrations, etc.

- Parameters and Acronyms

Instrument Parameters and Acronyms

Set point

PLF 		purge line flow			10 lpm

SF		sample flow				1.25 lpm

NOf		nitric oxide flow			3 mlpm

O3f		uv ozone monitor			0.5 lpm

Nafionf	Nafion flow				2.5 lpm

PLP		purge line pressure			~

NFP		Nafion pressure			~

PMT		Photomultiplier tube			Temp = -17 to -17.7 degC

RCT		Reaction Chamber Temperature	25 degC

RCP		Reaction chamber Pressure		18 Torr

HV		High Voltage				1500 V

Purge line length					23 m

DOY		Day of year

- Cruise Description

Below observations and events for this experiment are listed by day of year in 2007 and GMT time.

Day 191	2300: Ronald H. Brown left Galveston, TX; beginning of the cruise

Day 193	0300 – 1100: power loss, no data during that time

1500 – 1700: calibration from the mast

Day 195	1200 – 1500: cooler off, PMT temperature got up, data unusable

Day 196 	1500 – 1700: calibration from the mast

Day 197	1200: count drop in the FRCI signal which is not related to a drop in the 
		ozone concentration

Day 201	lightning storm in the night; instrument down during the night; fuse 
	broke; problems with data transfer from FRCI computer to NOAA 
computer where ozone data is acquisitioned to sonic data

Day 202	try to solve data transfer problem

to 		no continuous data for these days

Day 205	1700: instrument back to work, data transfer is working again

Day 206	PMT temperature was lower and RV temperature higher than before

1500 – 1700: calibration from the mast

Day 207	Lab View VI stopped; system needed to be reconfigured

Day 209	Hour 14 to 23 missing as ship had to go into Boston harbor due to ship 
compressor problems;

1700 – 1900: calibration from the box

Day 210	0000: pump of the Nafion® cycle broke

0001: pump of the Nafion® cycle replaced

Day 212	2200 – 0000: Calibration from the box

Day 214	1500 – 1700: Calibration from the mast

Day 216	1500 – 1700: Calibration from the box

1800: arrival in Boston; end of cruise

The fine filter and drying agent of the Nafion® dryer were changed every day at 0000 hour. Lag-time tests were carried out at the same time, usually before and after the filter change. During day 198 and 207 no lag-time tests were recorded.

- Calibrations

Calibrations were performed 8 times. Calibrations on day 209, 212, and 216 were performed from the box. The difference between calibration from the box and calibration from the mast is described in “Ozone flux measurements during the Gulf of Mexico and East Coast Carbon Cruise 2007”.

The average of the slope is then 2823.13 counts s-1, for the y-intercept it would be 2155.4 counts s-1. Therefore ozone in ppb would be

counts = (3823.13 ± 90.5)*[O3(ppb)]   +   (2155.4 ( 2852.1) 		(1)

counts = (4696.4 ± 96.5)*[O3(ppb)]   -   (2299.4 ( 3775)			(2)

Table 1

| Julian Day | slope | y-intercept |
| --- | --- | --- |
| 193 | 3785.8073 | 7619.643 |
| 196 | 3855.0367 | 2452.139 |
| 200 | 3641.2453 | 1733.1 |
| 206 | 4696.4051 | -2299.53 |
| 209 | 3876.8582 | -811.04 |
| 212 | 3917.1877 | 32.49338 |
| 214 | 3819.3459 | 3665.432 |
| 216 | 3866.4139 | 396.3371 |
| stdev | 90.475092 | 2852.056 |
| average | 3823.1279 | 2155.444 |

Day 206 is excluded from equation (1) due to an incorrect configuration of Measurements & Automatisation in LabView which affected the temperatures and thus the fast signal readings. As a result day 206 shows a significant departure from the other calibrations. The calibration of day 206 (equation (2)) is used for data from day 205 hour 17 to day 207 hour, which was the time period of the incorrect LabView configuration.

- Lag-time

The lag-time is described in the instruction files. On GOMECC, lag-time tests were performed at the 1st and 3rd third of the cruise. There, lag-time test were performed daily, before and after the filter change. The result was 8.8 s which is thus used in the data analysis for GOMECC.

- Figures

4.1 Concentrations along the ship track

To plot the mixing ratio or the flux along the ship track:

First use ‘map’ until delx=.25;

Then apply following:

CLIP  = 'on';

[x,y] = m_ll2xy(lon,lat,'clip',CLIP); colormap;

hold on ; scatter(x, y, 25, H)

H is the variable you want to plot. Usually you also want to apply a filter, thus you have to write x(p), … p is the filter, the data you want.

4.2 Wind direction along the ship track

To plot the wind direction you need the wind vectors U and V from the ‘da-file’, U or Ue is column 40 and V or Un column 39. Then do following:

use ‘map’ until delx=.25;	Then apply from ‘m_track’ file:

CLIP  = 'on';

[x,y] = m_ll2xy(lon,lat,'clip',CLIP); colormap;

hold on ; quiver(x, y, Ue, Un, 25)

- Filters

- Wind speed < 0.5 m*s-1

ii=find(u<0.5);

- Relative wind direction

ii=find(reldir<270 & reldir>90);

- Rain

ii=find(rain>5);

- Ozone mixing ratio below 3ppb or above 80 ppb since these values are usually not observed

ii=find(O3ppb<3 | O3ppb>80);

- Ozone standard deviation

ii=find(O3std>3 | O3std<-3);

- Delta Ozone; difference between 2 10min mean values. Is difference too big it indicates an unusual drop/increase

ii=find(deltaO3>3 | deltaO3<-3);

- motion correction

ii=find(sp2>2);

- Ship speed, indicates maneuver

ii=find(sig_u> 0.5);

- Standard deviation heading, indicates ship maneuver

ii=find(sig_h> 8);

- Filter developed by hand for calibration, wrong flow rates, missing data,…

ii=find(isnan(hand));

All filters are combined in the text-file ‘AMMA_filter_final’. The columns are:

jd=data(:,1);               %julianday

sp=data(:,2);               %ship speed

u=data(:,3);                %wind speed

reldir=data(:,4);           %relative wind direction

rain=data(:,5);             %NaN=rain>5

lag=data(:,6);              %lagtime

O3ppb_o=data(:,7);          %O3ppb

deltaO3_o=data(:,8);        %delta between 2 consecutive O3 values

O3std_o=data(:,9);          %standard deviation

O3ppb=data(:,10);           %O3ppb

deltaO3=data(:,11);         %delta between 2 consecutive O3 values

O3std=data(:,12);           %standard deviation

sp2=data(:,13);             %motion correction

sig_u=data(:,14);           %std dev of ship speed, m/s

sig_h=data(:,15);           %std dev of heading

hand=data(:,16);            %missing data, calibrations, flows,…

total_o=data(:,17);         %total filter crosscorr without lag

total=data(:,18);           %total filter constant lagtime without lag

‘_o’ indicates that this data is from cross correlation, equivalent other data is of constant lag-time method