## Sheet1
MotionPak Calibration Check, March 2018, pre-PISTON |  |  |  |  | g Boulder m/s^2
Test acceleration of gravity for each axis, in both directions |  |  |  |  | 9.7961 | Zumberge & Faller, JGR 1983
PSD Axis | V bow dn | V bow up | Cal Const | g bow | g stern | New Cal Const | Offset | g bow * | g stern *
Xacc | -3.787 | 3.759 | 3.7756 | 9.8296 | 9.7569 | 3.7745 | -0.028 | 9.7597 | 9.8324
| V port dn | V stbd | Cal Const | g port | g stbd | New Cal Const | Offset | g port * | g stbd *
Yacc | -3.849 | 3.7922 | 3.8215 | 9.8705 | 9.7249 | 3.8221 | -0.0568 | 9.7233 | 9.8689
| V normal | V invert | Cal Const | g up | g down | New Cal Const | Offset | g up * | g down *
Zacc | -3.682 | 3.714 | 3.7005 | 9.751 | 9.8358 | 3.6995 | 0.032 | 9.8385 | 9.7537
Problem with Yacc, port down. Voltage sometimes drops from -3.8 to -3.2. Ground problem? | [7 empty] | * applying offset and new cal const
(This is Xacc, port down in the MotPak coordinate system.)