NOAA/PSD  Ship-based Turbulent Flux Measurement System. 

Second Unit

# Rev 4.6

Turbulence Flux System setup on R/V Knorr.

Turbulence Portable Flux System setup on R/V Atlantis.

Turbulence Portable Flux System setup on R/V Atlantis.

Turbulence Portable Flux System setup on R/V Kilo Moana.

Turbulence Portable Flux System setup on R/V Kilo Moana.

Turbulence Flux System setup on main mast of R/V Ronald Brown.

The PSD Turbulent flux system consists of four components.

- A fast turbulence system with ship motion corrections mounted on the jackstaff.

The sensors are:

- Ultrasonic anemometer, Gill Wind Master Pro model #1561-PK-020.

- Motion-Pak, a  six axis motion detector built in house with Systron and Donner sensors.

- Solar and IR radiation sensors are:

- Radiometers are:

- Two Precision Spectral Pyranometer (PSP)from Appley Inc. One of them installed on the stabilized platform and one installed on a fix platform.

- Two Precision Infrared Radiometer (PIR) from Appley Inc. One of them installed on the stabilized platform and one installed on a fix platform.

- One Precision Pyranometer (CM22) from Kipp& Zonen. This one installed on the stabilized platform.

- One Pyrgeometer (CG4) from Kipp & Zonen. This one installed on the stabilized platform.

Radiometers are mounted on the highest and unobstructed sky-see location of vessel.

- Bulk Meteorology-sensor are:

- Mean T/RH sensor in aspirator, sensor model # HMP235 by Vaisala.

- Skin Surface Temperature (SST).SST measurement is made with a floating (YSI 46040) thermistor deployed off port side with outrigger.

- Optical Scientific Inc.-Optical Precipitation Sensor model ORG-815 DA.

- Li-COR 7500 fast CO2/H2O gas analyzer - mounted on the jack staff.

- Two differential GPS with pitch, roll and heading information.

| Table 1. Instruments and measurements for air-sea interaction, cloud, and precipitation studies. |
| --- |
| Item | System | Measurement |
| 1 | Air-sea flux system | Motion corrected turbulent fluxes |
| 2 | Pyranometer & Pyrgeometer | Downward solar radiative, IR flux |
| 3 | Bulk meteorology | Surface Water Temp, Air Temp, RH, Wind Speed |
| 4 | Optical Rain gauges | Rain falling Rate |
| 5 | Open Path LI-7500 Analyzer | CO2/H2O |

Data set of these instruments are logged, archived and displayed by two computers, with one computer used solely as the data acquisition system while the other computer serves as real time display data and as the platform for daily diagnostics.

All instruments data are transmitted to the acquisition computer via a selected serial communication port or via a selected wireless radio modem.

Acquisition System Structure

The acquisition System consist of three CR1000 data loggers, five self contained instruments and two Microsoft Windows XP professional computer.

- Campbell CR1000 data logger, slow mean data MEAN1(1Hz)

- Campbell CR1000 data logger, slow mean data MEAN2(1Hz)

- Campbell CR1000 data logger, slow mean data MEAN3(1Hz)

- Campbell CR1000 data logger, fast data Ship Motion(10Hz)

- Wind Master Pro Ultrasonic Anemometer Gill (10Hz)

- Open Path CO2/H2O Analyzer- LI-COR 7500 (10Hz)

- Two Hemisphere GPSVS100 systems (20Hz and 1Hz)

- Distance Meter Riegl LD90 3100VHS-FLP (10Hz)

- PS3DAS03 Microsoft Windows XP professional computer.

- PS3DAS04 Microsoft Windows XP professional computer.

PS3DAS03 computer collect real time data thru 16 serial ports interface PCI device using ports COM3 thru COM10 for flux data acquisition instruments.

COM11 thru COM18 are available for others instruments to be included on the present study.

Data Files

- All archived data file name is format as follows:
aaabyyjjjhh_raw.txt

aaa: Instrument identification

b:    Index , default is 0 for one instrument

YY:  Last two digits of year

JJJ:   julian day

HH:  hour

- Time stamp structure is as follows:

Timestamp: mmssmmm

mm: minutes

ss: seconds

mmm: milliseconds

- The data file folder structure is as follows:

D:\Data

Files created in “D:\Data” folder are move to D:\Data\<YYJJJ> at the end of the day (UTC time).

Archived files have the following format:

- Combined Ship Data - scs0YYJJJhh_raw.txt   (optional depending on ship data logging system)

- Ship Motion data – mot0YYJJJhh_raw.txt

- Met Data logger 1 – me10YYJJJhh_raw.txt

- Met Data logger 2 – me20YYJJJhh_raw.txt

- Met Data logger3  - me30YYJJJhh_raw.txt

- 7500 Licor gas analyzer – lic0YYJJJhh_raw.txt

- GPS Crescent VS100  – gps0YYJJJhh_raw.txt

- Ship Heading and Pitch  Crescent VS100– hed0YYJJJhh_raw.txt

- Ship Position Crescent  VS100– pos0YYJJJhh_raw.txt

- Ship Heading and Roll Crescent  VS100 – rol0YYJJJhh_raw.txt

- Sonic anemometer – son0YYJJJhh_raw.txt

- Wave Meter – lsr0YYJJJhh_raw.txt

Channel assignments are as follows on the data acquisition system:

- Serial Port COM1 received GPS data for computer time sync

- Serial Port COM2 reserved.

- Serial Port COM3-Gill Sonic Anemometer.

Data is acquired at 19200 baud rate and a sampling rate of 10Hz.

file name format: son0yyjjjhh_raw.txt

Data parameters are separate by a comma (,)

See bellow an example of the first two lines of an hourly Anemometer raw data file.

08/27/2007 18:00 Generic with ASCII Output (SN: 070514)

2545847 B,-001.20,-000.62,+000.22,M,+032.99,00,25

- Timestamp

- Node Address

- U

- V

- W

- M metric/sec units

- Sonic Temperature (degC)

- Status

- Serial Port COM4- LI-7500 Gas Analyzer.

Data is acquired at 19200 baud rate and a sampling rate of 10Hz.

file name format: lic0yyjjjhh_raw.txt

Data parameters are separate by a space.

See bellow an example of the first two lines of an hourly Licor7500 raw data file.

08/27/2007 18:00 Generic with ASCII Output (SN: 0772)

2540164 249	11.8338	446.337	36.53	83.3

- Timestamp

- Diag

- CO2

- H2O

- Temperature

- Pressure

'Diag' is a diagnostic, 'CO2' and 'H2O' are in units of mmol / m3, Temperature is in C˚, and Pressure is in kPa.

- Serial Port COM5- Mean1data logger-Temperature, Relative Humidity, Surface Water Temperature, and Rain Rate.

Data is acquired at 9600 baud rate and a sampling rate of 1Hz.

file name format: me10yyjjjhh_raw.txt

Data parameters are separate by a comma (,)

See bellow an example of the first two lines of an hourly Mean1 raw data file.

09/13/2007 20:00 Mean 1 Datalogger (SN: 7911)

0050811 13.11, 27.03, 0.214, 0.961, 2.258, 4.896, 0.056, 27.52, 18.42, 45.17, 0, 4.998, NAN, 441.2, 0.001,0,0,0.001, 0.01, 0,0,0.046,0.028,0.193,0,0,NAN,0

- Timestamp

- Data logger Battery voltage -average

- Data logger Temperature in Celsius -average

- Sea snake 2 wire resistance in ohms -average

- Vaisala Temperature voltage -average

- Vaisala Relative Humidity voltage -average

- ORG Rain Gauge Carrier Voltage -average

- ORG Rain Gauge Sensing Voltage -average

- Sea snake 2 wire temperature in Celsius -average

- Vaisala temp in Celsius -average

- Vaisala Relative Humidity in % -average

- ORG Rain Gauge sensing in mm/h -average

- Aspirator air flow >0 flow is on -average

- Sea snake 4 wire resistance in ohms -average

- Sea snake 4 wire Temperature in C –average

- Data logger Battery voltage –standard deviation

- Data logger Temperature in Celsius –standard deviation

- Sea snake 2 wire resistance in ohms –standard deviation

- Vaisala Temperature voltage –standard deviation

- Vaisala Relative Humidity voltage –standard deviation

- ORG Rain Gauge Carrier Voltage –standard deviation

- ORG Rain Gauge Sensing Voltage –standard deviation

- Sea snake 2 wire temperature in Celsius –standard deviation

- Vaisala temp in Celsius –standard deviation

- Vaisala Relative Humidity in % –standard deviation

- ORG Rain Gauge sensing in mm/h –standard deviation

- Aspirator air flow >0 flow is on –standard deviation

- Sea snake 4 wire resistance in ohms –standard deviation

- Sea snake 4 wire Temperature in C –standard deviation

Note : Usually  only one seasnake will be used to measure water surface temperature.

- Serial Port COM6-Mean 2 data logger- Solar Radiation PIR and PSP, Surface Water Temperature.

Data is acquired at 9600 baud rate and a sampling rate of 1Hz.

file name format: me20yyjjjhh_raw.txt

Data parameters are separate by a comma (,)

See bellow an example of the first two lines of an hourly Mean2 raw data file.

09/13/2007 20:00 Mean 2 Datalogger (SN: 7908)

0056349 12.69, 26.41, 0.507, 0.514, 0.452, 2.174,24.33, 23.63,114.2,246.5,2404, 30.11,NAN,491.7, 0.001, 0,0.009000001,0.004,0.021,0.231,0.905,0.397,5.353,26.16,6.461,0.063,NAN,0

1	Timestamp

2	Data logger Battery voltage -Average

3	Data logger Temperature -Average

4	Case_R-Average

5	Dome_R-Average

6	PIR voltage-Average

7	PSP voltage-Average

8	Case Temperature in C-Average

9	Dome temperature in C-Average

10	PIR Wm2-Average

11	PSP Wm2-Average

12	Sea snake 4 wire Resistance in ohms-Average

13	Sea snake 4 wire Temperature in C-Average

14	Sea temp 2 wire Resistance in ohms-Average

15	Sea temp 2 wire in C-Average

16	Data logger Battery voltage -Standard Deviation

17	Data logger Temperature  -Standard Deviation

18	Case_R -Standard Deviation

19	Dome_R -Standard Deviation

20	PIR voltage -Standard Deviation

21	PSP voltage -Standard Deviation

22	Case Temperature in C -Standard Deviation

23	Dome temperature in C -Standard Deviation

24	PIR Wm2 -Standard Deviation

25	PSP Wm2 -Standard Deviation

26	Sea snake 4 wire Resistance in ohms -Standard Deviation

27	Sea snake 4 wire Temperature in C -Standard Deviation

28	Sea temp 2 wire Resistance in ohms -Standard Deviation

29	Sea temp 2 wire in C -Standard Deviation

Note : Usually  only one seasnake will be used to measure water surface temperature.

- Serial Port COM7- All Ship instruments data
Data acquisition at 9600 baud rate with a sampling rate of 1Hz.

file name format: scs0yyjjjhh_raw.txt

1 	Time stamp
2 	To be determent

- Port COM8-Ship motion pitch, roll and yaw motion, measured with

Motion-Pak.

Data is acquired at 19200 baud rate and a sampling rate of 10Hz.

file name format: mot0yyjjjhh_raw.txt

Data parameters are separate by a comma (,)
See bellow an example of the first two lines of an hourly Motion-Pak  raw data file.

08/27/2007 18:00 Generic with ASCII Output (SN: 7459)

2535118 -0.03274211,0.9426998,0.001364254,-0.05184167,-0.1364254,-3.847198

- Timestamp

- Rate-x

- Rate-y

- Rate-z

- Acceleration-x

- Acceleration-y

- Acceleration-z

Rates and Accelerations are in Volts.

- Serial Port COM9- Ship Heading and Pitch with Crescent VS100

Data is acquired at 19200 baud rate and a sampling rate of 10Hz.

file name format: hed0yyjjjhh_raw.txt

Data parameters are separate by a comma (,)

See bellow an example of the first two lines of an hourly Ship heading and pitch raw data file.

07/09/2009 05:00 Generic with ASCII Output (SN: B1)

0000061 $PSAT,HPR,050001.20,250.45,0.89,,N*3D

1	Timestamp

2	$PSAT,HPR  identifier

3	time UTC time (HHMMSS.SS)

4	heading in degrees

5	pitch in degrees

6	roll in degrees

7	type N for GPS, G for gyro

Important note: This port is configured for Heading and Pitch,  therefore line 6 will show nothing.

- Serial Port COM10- Ship position GPS Cresent VS100

Data is acquired at 19200 baud rate and a sampling rate of 10Hz.

There are two kind of GPS file format that will be seen on GPS files, with 3 different GPS codes ($GPRMC), ($GPGGA), ($GPVTG)

file name format: gps0yyjjjhh_raw.txt

Data parameters are separate by a comma (,)

See bellow an example of the first two lines of an hourly Ship position GPS  raw data file. 06/12/2009 06:00 Generic with ASCII Output (A2)

0000077 $GPRMC,060000.40,A,3251.15194534, N,07956.64350164,W, 0.01,207.50, 120609, 7.3,W,D*3E

0000671 $GPVTG,145.12,T,152.46,M,0.01,N,0.02,K,D*22

$GPRMC coding

- Timestamp

- GPS code $GPRMC

- GPS time (HHMMSS.SS)

- Status ('A' = valid, 'V' = warning)

- Latitude in degrees, minutes and decimal minutes (DDMM.MMMMM)

- N or S latitude

- Longitude in degrees, minutes and decimal minutes(DDDMM.MMMMM)

- W or E longitude

- Speed-over-ground Z.Z in knots

- Track made good, reference to true north

- UTC date DDMMYY

- Magnetic Variation in degrees D.D

- Variation Sense v=E ;V=W

- *CC check sum

$GPVTG coding

- Timestamp

- GPS code $GPVTG

- T

- True course over ground, 000 to 359 degrees

- Magnetic course over ground, 000 to 359 degrees

- M

- Speed over ground, 00.0 to 99.9 knots

- N

- Speed over ground, 00.0 to 99.9 kilometers per hour

- K

- Check sum

- Serial Port COM11- Ship Heading and Roll with Crescent VS100

Data is acquired at 19200 baud rate and a sampling rate of 10Hz.

file name format: rol0yyjjjhh_raw.txt

Data parameters are separate by a comma (,)

See bellow an example of the first two lines of an hourly Ship heading and roll raw data file.

07/09/2009 08:00 Generic with ASCII Output (SN: B2)

0000061 $PSAT,HPR,080001.20,187.19,,10.53,N*06

1	Timestamp

2	$PSAT,HPR  identifier

3	time UTC time (HHMMSS.SS)

4	heading in degrees

5	pitch in degrees

6	roll in degrees

7	type N for GPS, G for gyro

Important note: This port is configured for heading and roll, therefore line 5 will show nothing.

- Serial Port COM12- Not in use

- Serial Port COM13 - Wave Measurement RIEGL

This setup is optional

Data are acquired at 9600 baud rate and a sampling rate of 20 Hz.

File name format lsr0yyjjjhh_raw.txt

Data parameters are separate by a comma (;)

See bellow an example of the first two lines of an hourly Wave measurement raw data file.

10/15/2008 01:00 LD90-3100VHS-FLP (SN: 12)

0000017 r7.389;a37;q80

- timestamp

- range in meters

- amplitud signal strength

- measurement quality

- Serial Port COM14 - Mean 3 data logger- Solar Radiation PIR and PSP, Surface Water Temperature.

Data is acquired at 9600 baud rate and a sampling rate of 1Hz.

file name format: me20yyjjjhh_raw.txt

Data parameters are separate by a comma (,)

See bellow an example of the first two lines of an hourly Mean3 raw data file.

09/13/2007 20:00 Mean 2 Datalogger (SN: 7908)

0056349 NAN, NAN, NAN, NAN, NAN, NAN, NAN, NAN, NAN, NAN, NAN, NAN, NAN

1	Timestamp

2	Data logger Battery voltage -Average

3	Data logger Temperature -Average

4	Case1_Rat-Average

5	Dome1_Rat-Average

6	Case1_R-Average

7	Dome1_R-Average

8	PIR1_V-Average

9	PSP1_V-Average

10	Case1_Temp-Average

11	Dome1_Temp-Average

12	PIR1_Wm2 -Average

13	PSP1_ Wm2 -Average

Data Processing Computers

Data Processing Computers tasks are:

- Computer name: PS03DAS03 Data Archive.

Computer Description: OptiPlex 330 Minitower, Pentium Dual Core E2180/2.0GHz,1M     800FSB

Main Software:

Data Acquisition Setup software, this is a windows XP service that allows to setup RS232 serial ports, set up archiving location, folder and files archiving names and archive the data with time stamp.

File Manager software, this is a Windows XP service allows to manage files as they are been archived.

Clock synchronization Windows XP service allows sync computer clock with GPS or Network.

- Computer name: PS03DAS04 Data Analysis.

Computer description: OptiPlex 330 Minitower, Pentium Dual Core  E2180/2.0GHz,1M     800FSB

File Manager software, this is a Windows XP service allows to manage files as they are been archived.

MathWorks used as mathematical computation, analysis, visualization, and algorithm developmental programs.

Labview is s a graphical programming environment used to display real time data.

Wireless communication Devices Setup

Six pairs of wireless radio modems are setup to send/receive data between instruments and acquisition computer. They are Sonic1 anemometer, Licor gas analyzer, Motion Pak ship motion, Mean1 data logger, Mean2 Data logger, and Riegl Distance Sensor. All modems are paired and configured on a point to point interface.

Two kind of radio modems are used:

- Data-Linc Group model PLantLinc 5000 with a transmitter frequency of 900 MHz. , refer to User guide for PLR5000.  These modems are used with the Mean1 data logger.

- Microhard System model Spectra 2420 with a transmitter frequency of 2.4 GHz., refer to Operating manual for Spectra 2420. These modems are used with the Sonic, Licor, Mean2, Motion Pak, and Laser.

Radio modem assignments are as follows:

| Parameter | Radio1 | Radio2 | Radio3 | Radio4 | Radio5 | Radio6 |
| --- | --- | --- | --- | --- | --- | --- |
| System | System 2 | System 2 | System 2 | System 2 | System 2 | System 2 |
| Serial # | 010-0014623 | 010-0014628 | 010-00115084 | 010-0015100 | 010-0015087 | 010-0015106 |
| S101 | 0 | 2 | 2 | 0 | 2 | 0 |
| S102 | 19200 | 19200 | 19200 | 19200 | 19200 | 19200 |
| S103 | 230400 | 230400 | 230400 | 230400 | 230400 | 230400 |
| S104 | 10 | 10 | 11 | 11 | 12 | 12 |
| S105 | 1 | 10 | 2 | 1 | 2 | 1 |
| S107 | spyder | spyder | spyder1 | spyder1 | spyder2 | spyder2 |
| S112 | 120 | 120 | 255 | 255 | 255 | 255 |
| S113 | 1 | 1 | 5 | 5 | 5 | 5 |
| S116 | 5 | 5 | 10 | 10 | 10 | 10 |
| S118 | 1 | 1 | 1 | 1 | 1 | 1 |
| S133 | 1 | 1 | 1 | 1 | 1 | 1 |
| S140 | 10 | 1 | 1 | 2 | 1 | 2 |
| Frequency | 2.4GHz | 2.4GHz | 2.4GHz | 2.4GHz | 2.4GHz | 2.4GHz |
| Connected to | COMM8 of PC | CR1000 s/n 7459 for Motion Pak | COMM3 of PC | Sonic Interface | COMM4 of PC | Licor Interface |
| Restrictions | none | None | none | none | none | None |

| Parameter | Radio7 | Radio8 | Radio9 | Radio10 | Radio11 | Radio12 |
| --- | --- | --- | --- | --- | --- | --- |
| System | System 2 | System 2 | System 2 | System 2 | System 2 | System 2 |
| Serial # | 010-0015098 | 010-0015092 | 010-0015085 | 010-0015082 | C090013 | C090012 |
| S101 | 0 | 2 | 0 | 2 | Master=0 | Slave=1 |
| S102 | 9600 | 9600 | 9600 | 9600 | 9600 | 9600 |
| S103 | 230400 | 230400 | 230400 | 230400 | n/a | n/a |
| S104 | 13 | 13 | 14 | 14 | 255 | 255 |
| S105 | 1 | 2 | 1 | 2 | 885-1137 | 885-1200 |
| S107 | spyder3 | spyde3 | spyder4 | spyder4 | n/a | n/a |
| S112 | 255 | 255 | 255 | 255 |  |  |
| S113 | 5 | 5 | 5 | 5 |  |  |
| S116 | 10 | 10 | 10 | 10 |  |  |
| S118 | 1 | 1 | 1 | 1 |  |  |
| S133 | 1 | 1 | 1 | 1 |  |  |
| S140 | 2 | 1 | 2 | 1 | 885-1200 | 885-1137 |
| Frequency | 2.4GHz | 2.4GHz | 2.4GHz | 2.4GHz | 902 MHz | 902 MHz |
| Connected to | COMM 13 of PC | Laser Interface | COMM6 of PC | CR1000 s/n 7908 for Mean 2 | COMM5 of PC | CR1000 s/n 7911 for Mean 1 |
| Restrictions | none | None | none | none | none | None |

NOTE: Radio Modem parameters are configured before installation;  if replacement is needed replace with a pair of radios and program them accordingly to manual.

Note: This figure represents a standard configuration- installation of the portable flux system; it does not necessarily reflect the actual configuration set up on the field.

Note: This figure represents a standard configuration- installation of the portable flux system; it does not necessarily reflect the actual configuration set up on the field.

Sonic Interface Diagram

7500 Licor Interface Diagram

Mean 1 Interface Diagram

Mean 2 Interface Diagram

Motion Pak and Sprayer Interface Diagram

Interface Diagram for T/RH, Licor, Laser

Optical Rain Gauge interface Diagram

Datalogger programs

-Motion Pak Program

'CR1000 Series Datalogger

‘Program Name: Second_MotoinPak_1.CR1

'Created by Sergio Pezoa

'Date: Feb 15, 2007

'Declare Public Variables

'offsets as 3/31/09 -0.00955,0.00136,-0.01501,-0.05321,0.00819,-3.84311

' offset have been changed as 2/4/2010

' Sprayer control- addition - modified on 7/29/2010 by SP

PipeLineMode

Public Meas(7)

Public PortOn

Alias meas(1) = xrate

Alias meas(2) = yrate

Alias meas(3) = zrate

Alias meas(4) = xaccel

Alias meas(5) = yaccel

Alias meas(6) = zaccel

Public OutString As String * 1000

'Define Data Tables

DataTable (Motion,1,-1)

Sample (6,Meas(),FP2)

EndTable

'Main Program

BeginProg

'SerialOpen (ComSDC7,19200,0,0,2000)

SerialOpen (ComRS232,19200,0,0,200)

Scan (100,mSec,800,0)

'Sprayer controller

If TimeIntoInterval(0,30,Sec) Then PortOn=0 'set Port 1 to 0

PortSet(9, PortOn)   'Set port 1 low for time interval

If TimeIntoInterval(1,43170,Sec) Then PortOn=1 'set PortOn to 0

PortSet(9, PortOn)   'Set port 1 high for time interval

'motion pak voltages reading

VoltSe (Meas(1),1,mV5000,1,1,0,250,0.001,0.006)'xrate offset = 0.0109  10/18/10

VoltSe (Meas(2),1,mV5000,2,1,0,250,0.001,0.001) 'yrate offset = -0.0027  10/18/10

VoltSe (Meas(3),1,mV5000,3,1,0,250,0.001,0.013) 'zrate offset = 0.0156   10/18/10

VoltSe (Meas(4),1,mV5000,4,1,0,250,0.001,-0.016) 'xaccel offset =0.008  10/18/10

VoltSe (Meas(5),1,mV5000,5,1,0,250,0.001,-0.012) 'yaccel offset = 0  10/18/10

VoltSe (Meas(6),1,mV5000,6,1,0,250,0.001,0)'no zaccel offset   10/18/10

OutString = xrate+CHR(44)+yrate+CHR(44)+zrate+CHR(44)+xaccel+CHR(44)+yaccel+CHR(44)+zaccel+CHR(13)+CHR(10)

'SerialOut (ComSDC7,OutString,"",0,0)

SerialOut (ComRS232,OutString,"",0,0)

SerialFlush (COMRS232)

CallTable Motion

NextScan

EndProg

-Mean 1 program

'Program name: T_RH_RAIN_SEAT_ASPIR_VER7.CR1

'By Sergio Pezoa

'Mean1

'Log sea temp 2 wire,vaisala T/RH, Rain Gauge, aspirator and sea Temp 4 wire

'Date written: 8/30/2007

'Modified on Jan/26/2011

'IF "NAN" make it -9999

Public Flag(1) as boolean

public outputflag

Public OutString1 as string * 100

Public OutString2 as string * 100

Public OutString3 as string * 100

Public OutString4 as string * 100

Public TestRecord(28)

public count

'Declared variables and units

public Batt_v, PTemp

Public Seatemp_r

Public Seatemp

Public Snake_r4W

Public vaisala_v(2)

Public ORG_Sig_v

Public ORG_Car_v

Public SeatempC

Public SnakeC_4W

Public vaisT_C

Public vaisRH

Public ORG_mmhr

Public airflow

DIM  I

Alias vaisala_v(1)=vaisT_v

Alias vaisala_v(2)=vaisRH_v

'Declare Constants

Const c1=0.0010295

Const c2=0.0002391

Const c3=0.0000001568

'Constants for ysi serie 46040  c4,c5,c6 -5 to 40 C

Const c4=0.001399937

Const c5=0.000237854

Const c6=0.000000097

'Contant wire resistance for seasnake

Const r_cond=5.120

Const high=true

Const low=false

'Define Data Table

DataTable(Mean1,outputflag,-1)

DataInterval(0,1,Min,-1)

'average

Average(1,Batt_v,FP2,0)

Average (1,PTemp,FP2,0)

Average(1,Seatemp_r,FP2,0)

Average(1,vaisT_v,FP2,0)

Average(1,vaisRH_v,FP2,0)

Average(1,ORG_Car_v,FP2,0)

Average(1,ORG_Sig_v,FP2,0)

Average(1,SeatempC,FP2,0)

Average(1,vaisT_C,FP2,0)

Average(1,vaisRH,FP2,0)

Average(1,ORG_mmhr,FP2,0)

Average(1, airflow, FP2,0)

Average (1,Snake_r4W,FP2,0)

Average (1,SnakeC_4W,FP2,0)

'standard dev

StdDev(1,Batt_v,FP2,0)

StdDev(1,PTemp,FP2,0)

StdDev(1,Seatemp_r,FP2,0)

StdDev(1,vaisT_v,FP2,0)

StdDev(1,vaisRH_v,FP2,0)

StdDev(1,ORG_Car_v,FP2,0)

StdDev(1,ORG_Sig_v,FP2,0)

StdDev(1,SeatempC,FP2,0)

StdDev(1,vaisT_C,FP2,0)

StdDev(1,vaisRH,FP2,0)

StdDev(1,ORG_mmhr,FP2,0)

StdDev(1,airflow,FP2,0)

StdDev(1,Snake_r4W,FP2,0)

StdDev(1,SnakeC_4W,FP2,0)

EndTable

'Main Program

BeginProg

SerialOpen (ComRS232,9600,0,0,10000)

Scan(1,Sec, 3, 0)

If TimeIntoInterval (0,60,Sec) then flag(1)=high

I=i+1

if i > 60

i=0

endif

count=I

'seasnakecheck

'1H & GND with 10K resistor between 1H and E1

BRHalf(Seatemp_r, 1,mV250, 1,Vx1, 1, 250, True, 0, _60Hz, 1, 0)

Seatemp = 10000*(Seatemp_r/(1.0-Seatemp_r))

IF StrComp(Seatemp, 1 ) < 0 then

SeatempC = -9999

ELSE

SeatempC = 3915.9/(r_cond+ Ln(Seatemp - 10))-273.16

ENDIF

'sea snake Half Bridge, 4 Wire measurements Seatemp_r4W:

'7H & 7L ,8H & 8L ,E1

'(wire1,2) ==/\/\/\/\/\== (wire3,4)

'CH7- wire1(yellow)to 7H/10k resistor between 7H and 7L/7L to E3

'wire3(green)to GND

'CH8- wire2(red)to 8H/wire4(blue)to 8L

BrHalf4W(Snake_r4W,1,mV2500,mV2500,7,3,1,2500,False,False,0,_60Hz,10031,0)

IF StrComp(Snake_r4W,"INF") = 0 then

SnakeC_4W=-9999

ELSE

SnakeC_4W = 1/(c4+c5*Ln(Snake_r4W)+c6*PWR(Ln(Snake_r4W),3))+(-273.16)

ENDIF

'Aspirator check

'2H , E2

BRHalf(airflow, 1,mV2500, 3,Vx2, 1,2000,False, 0,250,-5,5)

PanelTemp (PTemp,250)

Battery(Batt_v)

'vaisala temp nad rh voltages

'3H & 3L for temp 4H & 4L for humidity

VoltDiff(vaisala_v(), 2, mV5000, 3,False, 0, _60Hz, 0.001, 0)

vaisT_c = vaisT_v * 40 - 20

vaisRH = vaisRH_v * 20

'rain gauge carrige voltages

'5H &5L

VoltDiff(ORG_Car_v, 1, mV5000,5 ,False, 0, _60Hz, 0.001, 0)

'rain gauge sense voltage

'6H & 6L

VoltDiff(ORG_Sig_v, 1, mV5000, 6,False, 0, _60Hz, 0.001, 0)

ORG_mmhr = 0

'

If  ((ORG_Sig_v) >= 0.085) Then

'rain rate in mm/hr with an standard calibration (500 mm/hr)

ORG_mmhr = (ORG_Sig_v) ^ 2 * 20 - 0.05

EndIf

If TimeIntoInterval(0,60,Sec) then outputflag=high

If Flag(1)=true then

GetRecord (TestRecord(),Mean1,1)

OutString1=TestRecord(1)+CHR(44)+TestRecord(2)+CHR(44)+TestRecord(3)+CHR(44)+TestRecord(4)+CHR(44)+TestRecord(5)+CHR(44)+TestRecord(6)+CHR(44)+TestRecord(7)+CHR(44)

OutString2=TestRecord(8)+CHR(44)+TestRecord(9)+CHR(44)+TestRecord(10)+CHR(44)+TestRecord(11)+CHR(44)+TestRecord(12)+CHR(44)+TestRecord(13)+CHR(44)+TestRecord(14)+CHR(44)

OutString3=TestRecord(15)+CHR(44)+TestRecord(16)+CHR(44)+TestRecord(17)+CHR(44)+TestRecord(18)+CHR(44)+TestRecord(19)+CHR(44)+TestRecord(20)+CHR(44)+TestRecord(21)+CHR(44)

OutString4=TestRecord(22)+CHR(44)+TestRecord(23)+CHR(44)+TestRecord(24)+CHR(44)+TestRecord(25)+CHR(44)+TestRecord(26)+CHR(44)+TestRecord(27)+CHR(44)+TestRecord(28)+CHR(13)+CHR(10)

SerialOut (ComRS232,OutString1+OutString2+OutString3+OutString4,"",0,0)

Flag(1)=false

EndIf

CallTable Mean1

NextScan

EndProg

-Mean 2 program

'CR1000

'Program name: PSP_PIR_SEC_SYS_VER9.CR1

'Created by Sergio Pezoa

'Mean 2

'Updated 26/01/11

'Logs Sea snake with either 2 or 4 wires, PSP and PIR...

'Declare Variables and Units

'9/28/07 "NAN" changed to -9999

'04/01/09 added to second system ....name changed

'ONE pyranometer EPPLEY,ONE pyrgeometer EPPLEY, ONE pyranometer K & Z,ONE pyrgeometer  K & Z,

'Sea Snake, Atmosphoric Pressure

'updated pir's and psp's coeficients 26\01\2011

'channels used as:

'1H=Eppley PIR Case Temp 1

'1L=Eppley PIR Dome Temp 3

'2H=Eppley PIR thermopile 5

'2L=Eppley PIR thermopile 6

'----------

'3H=Eppley PSP thermopile

'3L=Eppley PSP thermopile

'----------

'4H=2W sea snake

'4L= to Ground

'----------

'5H= 10K between 5H&E1 K&Z PIR

'5L= THERMISTOR yellow K&Z PIR

'AG= THERMISTOR green K&Z PIR

'6H= +THERMOPILE PIR  K&Z PIR

'6L= -THERMOPILE PIR K&Z PIR

'-----------

'7H= +THERMOPILE K&Z PSP pin3

'7L= -THERMOPILE K&Z PSP pin1

'-----------

'8H= Pressure sensor PTB220 voltage

'8L= to Ground

'Declare Variables and Units

Public Batt_Volt,PTemp

Public Case_Rat(2)

Public PIRV(2)

Public PIRKZ_RR

Public PIRKZ_R

Public PIRKZ_V

Public PSPKZ_V

Public PIRKZ_Wm2

Public PIRKZ_C

Public PSPKZ_Wm2

Public Di_volt(3)

Public PIRWm2

Public PSPWm2

Public Case_R

Public Dome_R

Public CaseTemp

Public DomeTemp

Public Seatemp_r

Public Seatemp

Public SeatempC

Public pressure_v

Public hpa

Public outputflag

Public Flag (1) As Boolean

Public OutString As String * 100

Public OutString1 As String * 100

Public OutString2 As String * 100

Public OutString3 As String * 100

Public OutString4 As String * 100

Public OutString5 As String * 100

Public OutString6 As String * 100

Public count1,TestNum(40)

Public TimeValue As String * 25

Alias Case_Rat(1) = CaseRat

Alias Case_Rat(2) = DomeRat

Alias PIRV(1) = PIR_V

Alias PIRV(2) = PSP_V

Alias Di_volt(1)=Diff1

Alias Di_volt(2)=Diff2

Alias Di_volt(3)=Diff3

'Declare Other Variables

'Declare Constants

'PSP and PIR constants from document 10/23/97 requested by J Hare

' using YSI 44031

Const c1=0.0010295

Const	c2=0.0002391

Const	c3=0.0000001568

'Constants for YSI serie 46040  c4,c5,c6 do not change

Const c4=0.001399937

Const c5=0.000237854

Const c6=0.000000097

'Contant wire resistance for seasnake

Const r_cond=5.120

'PIR cte 3.84e-6 s/n 34302F3 Jan/12/2010

Const PIRCnst = 0.00384

'PSP cte 8.64e-6  s/n 34290F3 Jan/12/10

Const PSPCnst = 0.00864

Const high=true

Const low = false

'Coefficients fot YSI 44031 PIR1 0 to 40 C

' to calculate thermistor temp

Const c7=0.001017394

Const c8=0.000241046

Const c9=0.000000149

'PIR K&Z CG4 cte 10.59  s/n 50770 Jan/12/2011

Const PIRKZ_Cte = 0.01059

'PSP K&Z CM22 cte 8.99  s/n 50122 Jan/12/2011

Const PSPKZ_Cte = 0.00899

'Define Data Tables

DataTable(Metsec,outputflag,-1)

DataInterval(0,1,Min,-1)

'average

Average(1,Batt_Volt,FP2,0)

Average (1,PTemp,FP2,0)

Average (1,CaseRat,FP2,0)

Average (1,DomeRat,FP2,0)

Average (1,PIR_V,FP2,0)

Average (1,PSP_V,FP2,0)

Average (1,CaseTemp,FP2,0)

Average (1,DomeTemp,FP2,0)

Average (1,PIRWm2,FP2,0)

Average (1,PSPWm2,FP2,0)

Average(1,PIRKZ_R,FP2,0)

Average(1,PIRKZ_RR,FP2,0)

Average(1,PIRKZ_V,FP2,0)

Average(1,PSPKZ_V,FP2,0)

Average(1,PIRKZ_Wm2,FP2,0)

Average(1,PIRKZ_C,FP2,0)

Average(1,PSPKZ_Wm2,FP2,0)

Average (1,Seatemp_r,FP2,0)

Average(1,SeatempC,FP2,0)

Average(1,hpa,FP2,0)

'stabdard dev

StdDev(1,Batt_Volt,FP2,0)

StdDev (1,PTemp,FP2,0)

StdDev (1,CaseRat,FP2,0)

StdDev (1,DomeRat,FP2,0)

StdDev (1,PIR_V,FP2,0)

StdDev (1,PSP_V,FP2,0)

StdDev (1,CaseTemp,FP2,0)

StdDev (1,DomeTemp,FP2,0)

StdDev (1,PIRWm2,FP2,0)

StdDev (1,PSPWm2,FP2,0)

StdDev(1,PIRKZ_R,FP2,0)

StdDev(1,PIRKZ_RR,FP2,0)

StdDev(1,PIRKZ_V,FP2,0)

StdDev(1,PSPKZ_V,FP2,0)

StdDev(1,PIRKZ_Wm2,FP2,0)

StdDev(1,PIRKZ_C,FP2,0)

StdDev(1,PSPKZ_Wm2,FP2,0)

StdDev (1,Seatemp_r,FP2,0)

StdDev(1,SeatempC,FP2,0)

StdDev(1,hpa,FP2,0)

EndTable

'Main Program

BeginProg

SerialOpen (ComRS232,9600,0,0,10000)

Scan(1,Sec,0,0)

If TimeIntoInterval (0,60,Sec) Then Flag(1) = high

TimeValue=Metsec.Timestamp(1,1)

Battery(Batt_Volt)

PanelTemp(PTemp,250)

'Generic Half Bridge measurements Eppley PIR CaseR(1):

'1H case resistance with 10K resistor between 1H and E1

'1L dome resistance with 10K resistor between 1L and E1

BrHalf(Case_Rat(),2,Autorange,1,Vx1,2,500,False,0,250,1,0)

Case_R=10000*(CaseRat/(1-CaseRat))

Dome_R=10000*(DomeRat/(1-DomeRat))

CaseTemp = 1/(c1+c2*Ln(Case_R)+c3*PWR(Ln(Case_R),3))+(-273.16)

DomeTemp = 1/(c1+c2*Ln(Dome_R)+c3*PWR(Ln(Dome_R),3))+(-273.16)

'Generic Differential Voltage measurements Eppley PIR PIR_V:

'Diff 2H,2L

VoltDiff(PIR_V,1,mV25,2,False,0,_60Hz,1.0,0)

IF StrComp(PIR_V,"NAN")= 0 then

PIR_V = -9990

PIRWm2 = -9991

ELSE

PIRWm2=(1/PIRCnst)*PIR_V

ENDIF

'Generic Differential Voltage measurements Eppley PSP 3H, 3L

VoltDiff(PSP_V,1,mV7_5,3,False,0,_60Hz,1.0,0)

IF StrComp(PSP_V,"NAN")= 0 then

PSP_V = -9990

PSPWm2 = -9991

ELSE

PSPWm2=(1/PSPCnst)*PSP_V

ENDIF

'sea snake measuremnet check

'4H & GND with 10K resistor between 4H and E3

BRHalf(Seatemp_r, 1,mV250, 7,Vx3, 1, 250,False,0,250, 1,-0.014)

Seatemp = 9986*(Seatemp_r/(1.0-Seatemp_r))

if StrComp(Seatemp, 1 ) < 0 then

SeatempC = -9999

else

SeatempC = 3915.9/(r_cond+ Ln(Seatemp - 10))-273.16

endif

'Half bridge measurements K&Z PIR thermistor:

BrHalf (PIRKZ_R,1,mV2500,9,Vx2,1,500,False,0,250,1.0,0)

PIRKZ_RR = 10000*(PIRKZ_R/(1-PIRKZ_R))

IF StrComp(PIRKZ_RR,"NAN")= 0 then

PIRKZ_RR = -9990

PIRKZ_C = -9991

ELSE

'calculated thermistor temperature

PIRKZ_C =1/(c7+c8*Ln(PIRKZ_RR)+c9*PWR(Ln(PIRKZ_RR),3))+(-273.16)

ENDIF

'PIR K&Z Differential Voltage measurement PIRKZ_V() thermopile:

VoltDiff(PIRKZ_V,1,mV25,6,False,0,_60Hz,1.0,0)

IF StrComp(PIRKZ_V,"NAN")=0 then

PIRKZ_V= -9990

PIRKZ_Wm2= -9991

else

PIRKZ_Wm2=(1/PIRKZ_Cte)*PIRKZ_V

ENDIF

'PSP K&Z Differential Voltage measurement PSPKZ_V thermopile:

VoltDiff(PSPKZ_V,1,mV25,7,False,0,_60Hz,1.0,0)

IF StrComp(PSPKZ_V,"NAN")= 0 then

PSPKZ_V= -9990

PSPKZ_Wm2= -9991

else

PSPKZ_Wm2=(1/PSPKZ_Cte)*PSPKZ_V

ENDIF

'Differential Vaisala Presure measuremnet 8H,8L

VoltDiff (pressure_v,1,mV5000,8,False,0,_60Hz,0.001,0.013)

'VoltSe (pressure_v,1,mV5000,15,1,0,250,0.001,0.033)

hpa=pressure_v*120+500

'boulder height h=1655meters

'hpa= pressure+0.1176*h

'Call Data Tables and Store Data

If TimeIntoInterval(0,60,sec) then outputflag=high

If Flag (1)=true Then

GetRecord (TestNum(),Metsec,1)

OutString1 =TestNum(1)+CHR(44)+TestNum(2)+CHR(44)+TestNum(3)+CHR(44)+TestNum(4)+CHR(44)+TestNum(5)+CHR(44)+TestNum(6)+CHR(44)+TestNum(7)+CHR(44)

OutString2 =TestNum(8)+CHR(44)+TestNum(9)+CHR(44)+TestNum(10)+CHR(44)+TestNum(11)+CHR(44)+TestNum(12)+CHR(44)+TestNum(13)+CHR(44)+TestNum(14)+CHR(44)

OutString3 =TestNum(15)+CHR(44)+TestNum(16)+CHR(44)+TestNum(17)+CHR(44)+TestNum(18)+CHR(44)+TestNum(19)+CHR(44)+TestNum(20)+CHR(44)+TestNum(21)+CHR(44)

OutString4 =TestNum(22)+CHR(44)+TestNum(23)+CHR(44)+TestNum(24)+CHR(44)+TestNum(25)+CHR(44)+TestNum(26)+CHR(44)+TestNum(27)+CHR(44)+TestNum(28)+CHR(44)

OutString5 =TestNum(29)+CHR(44)+TestNum(30)+CHR(44)+TestNum(31)+CHR(44)+TestNum(32)+CHR(44)+TestNum(33)+CHR(44)+TestNum(34)+CHR(44)+TestNum(35)+CHR(44)

OutString6 =TestNum(36)+CHR(44)+TestNum(37)+CHR(44)+TestNum(38)+CHR(44)+TestNum(39)+CHR(44)+TestNum(40)+CHR(13)+CHR(10)

SerialOut (ComRS232,OutString1+OutString2+OutString3+OutString4+OutString5+OutString6,"",0,0)

Flag (1)=false

EndIf

CallTable(Metsec)

NextScan

EndProg

-Mean 3 program

'CR1000

'Created by Sergio Pezoa

'Feb/03/2011

'One pyranometer Eppley

'One pyrgeometer Eppley

'----------

'Channels used as:

'1H=Eppley PIR Case Temp 1 pin D

'GND Case pin E

'-----------

'1L=Eppley PIR Dome Temp 3 pin F

'GND Dome pin G

'-----------

'2H=Eppley PIR thermopile 5 pin C

'2L=Eppley PIR thermopile 6 pin A

'3H=Eppley PSP thermopile connector pin B

'3L=Eppley PSP thermopile connector pin A

'GND pin C with a jumper to A

'Declare Variables and Units

Public Batt_Volt,PTemp

Public Case1_Rat

Public Case1_R

Public Dome1_Rat

Public Dome1_R

Public Case1_Temp

Public Dome1_Temp

Public PIR1_V

Public PIR1_Wm2

Public PSP1_V

Public PSP1_Wm2

Public Outnum(12)

Public OutString1 As String * 100

Public OutString2 As String * 100

Public outputflag

Public Flag(1) As Boolean

'Declare Other Variables

Units Batt_Volt=Volts

'Declare Constants

Const high= true

Const low= false

'PSP and PIR constants requiested by Jeff Hare document 10/23/97

'Campbell Scientifics Inc. App. Note Code: 2RA-H

' using YSI 44031

Const c1=0.0010295

Const	c2=0.0002391

Const	c3=0.0000001568

'PIR cte 4.05e-6 s/n 30558F3 Feb/03/2011

Const PIRCnst = 0.00405

'PSP cte 8.56e-6  s/n 30593F3 Feb/03/2011

Const PSPCnst = 0.00856

'Define Data Tables

DataTable(PSP_PIR,outputflag,-1)

DataInterval(0,1,Min,-1)

'Averages

Average(1,Batt_Volt,FP2,0)

Average(1,PTemp,FP2,0)

Average(1,Case1_Rat,FP2,0)

Average(1,Dome1_Rat,FP2,0)

Average(1,Case1_R,IEEE4,0)

Average(1,Dome1_R,IEEE4,0)

Average(1,PIR1_V,FP2,0)

Average(1,PSP1_V,FP2,0)

Average(1,Case1_Temp,FP2,0)

Average(1,Dome1_Temp,FP2,0)

Average(1,PIR1_Wm2,FP2,0)

Average(1,PSP1_Wm2,FP2,0)

EndTable

'Main Program

BeginProg

SerialOpen (ComRS232,9600,0,100,10000)

Scan(1,Sec,0,0)

If TimeIntoInterval (0,60,sec) Then Flag(1) = high

'Default Datalogger Battery Voltage measurement Batt_Volt:

Battery(Batt_Volt)

PanelTemp(PTemp,250)

'Generic Half Bridge measurements Eppley PIR CaseR(1):

'1H case resistance with 10K resistor between 1H and E1

'1L dome resistance with 10K resistor between 1L and E1

BrHalf(Case1_Rat,1,mV2500,1,Vx1,2,500,False,0,250,1,0)

BrHalf(Dome1_Rat,1,mV2500,2,Vx1,2,500,False,0,250,1,0)

Case1_R=10000*(Case1_Rat/(1-Case1_Rat))

Dome1_R=10000*(Dome1_Rat/(1-Dome1_Rat))

Case1_Temp = 1/(c1+c2*LN(Case1_R)+c3*PWR(LN(Case1_R),3))+(-273.16)

Dome1_Temp = 1/(c1+c2*LN(Dome1_R)+c3*PWR(LN(Dome1_R),3))+(-273.16)

'Generic Differential Voltage measurements Eppley PIR PIR_V:

'Diff 2H,2L

VoltDiff(PIR1_V,1,mV25,2,False,0,_60Hz,1.0,0)

If StrComp(PIR1_V,"NAN")= 0 Then

PIR1_V = -9990

PIR1_Wm2 = -9991

Else

PIR1_Wm2=(1/PIRCnst)*PIR1_V

EndIf

'Generic Differential Voltage measurements Eppley PSP 3H, 3L

VoltDiff(PSP1_V,1,mV25,3,False,0,_60Hz,1.0,0)

If StrComp(PSP1_V,"NAN")= 0 Then

PSP1_V = -9990

PSP1_Wm2 = -9991

Else

PSP1_Wm2=(1/PSPCnst)*PSP1_V

EndIf

'Call Data Tables and Store Data

If TimeIntoInterval(0,60,sec) Then outputflag=high

If Flag(1)= true Then

GetRecord (Outnum(),PSP_PIR,1)

OutString1 = Outnum(1)+CHR(44)+Outnum(2)+CHR(44)+Outnum(3)+CHR(44)+Outnum(4)+CHR(44)+Outnum(5)+CHR(44)+Outnum(6)+CHR(44)

OutString2 = Outnum(7)+CHR(44)+Outnum(8)+CHR(44)+Outnum(9)+CHR(44)+Outnum(10)+CHR(44)+Outnum(11)+CHR(44)+Outnum(12)+CHR(13)+CHR(10)

SerialOut (ComRS232,OutString1+OutString2,"",0,100)	'Send data string over ComRS232

Flag (1)= false

EndIf

CallTable(PSP_PIR)

NextScan

EndProg

Coefficient Table of Sensors

| Sensor | Calibration coefficient | Make / Model | Serial Number | Date of calibration |
| --- | --- | --- | --- | --- |
| Precision Spectral Pyranometer | 0.00864 | Eppley / PSP | 34290F3 | Jan 12, 2010 |
| Precision Spectral Pyranometer | 0.00405 | Eppley / PSP | 30558F3 | Feb 3, 2011 |
| Precision Infrared Radiometer | 0.00384 | Eppley / PIR | 34302F3 | Jan 12, 2010 |
| Precision Infrared Radiometer | 0.00856 | Eppley / PIR | 30593F3 | Feb 3, 2011 |
| Precision Spectral Pyranometer | 0.00899 | Kipp & Zonen CM22 | 50122 | Jan 12, 2011 |
| Precision Infrared Radiometer | 0.01059 | Kipp & Zonen CG4 | 50770 | Jan 12, 2011 |
| Motion Pak ZERO xrate | 0.0109 | Systron & Donner |  |  |
| Motion Pak ZERO yrate | -0.0027 | Systron & Donner |  |  |
| Motion Pak ZERO zrate | 0.0156 | Systron & Donner |  |  |
| Motion Pak Zero xaccel | 0.008 | Systron & Donner |  |  |
| Motion Pak ZERO yaccel | 0 | Systron & Donner |  |  |
| Motion Pak ZERO zaccel | 0 | Systron & Donner |  |  |
| Sea Snake thermistor | C4=0.001399937 | YSI 46040 series |  | Dec 12, 2009 |
| Sea Snake thermistor | C5=0.00237854 | YSI 46040 series |  |  |
| Sea Snake thermistor | C6=0.000000097 | YSI 46040 series |  |  |

Data Analysis Computer

The Data Analysis computer produces daily graphs with the MATLAB Flux Evaluation routines compiled to run automatically.

Optional Lab-View routines can be access to create real time data display with the purpose to troubleshoot possible problems with all instruments. See example displays showed below

Hemisphere GPS VS100 set up summary

A GPS VS100 system can be setup to measure heading and pitch or heading and roll but not both.

To get heading and pitch measurements, two steps to follow are needed.

- Primary and secondary antenna need to be in line between the bow and the stern of the ship with the primary antenna on the bow side.

- The following command need to be send to the crescent unit

“$JATT,ROLL,NO <CR><LF>”      (default)

To get heading and roll measurements, two setups are needed

- Primary and secondary antenna need to be in line between the port and starboard of the ship with the primary antenna on the port side.

2.   The following command need to be send to the crescent unit

“$JATT,ROLL,YES <CR><LF>”

The “$JASC” command allows the GPS data messages on at a particular update

rate, to be turned on or off. When turning messages on, various update

rates are available, depending on what the requirements are.

This command has the following layout:

$JASC,MSG,R[,OTHER]<CR><LF>

Where “MSG” is the name of the data message and “R” is the message

rate, as shown in the table below. Sending the command without the

optional “,OTHER” data field (without the brackets) will enact a change

on the current port.

Sending a command with a zero value for the “R” field turns off a

message.

| MSG | R (Hz) | Description |
| --- | --- | --- |
| GPGBS 20, 10, 2, 1, 0 or .2 Satellite fault detection used for RAIM |
| GPGGA 20, 10, 2, 1, 0 or .2 GPS fix data |
| GPGLL 20, 10, 2, 1, 0 or .2 Geographic position -latitude/longitude |
| GPGNS 20, 10, 2, 1, 0 or .2 GNSS fix data |
| GPGRS 20, 10, 2, 1, 0 or .2 GNSS range residuals |
| GPGSA 1 or 0 GNSS (Global Navigation Satellite System DOP (and active satellites)) |
| GPGST 1 or 0 GNSS pseudorange error statistics |
| GPGSV 1 or 0 GNSS satellite in view |
| GPRMC 10, 2, 1, 0 or .2 Recommended minimum specific GNSS data |
| GPRRE 1 or 0 Range residual message |
| GPVTG 10, 2, 1, 0 or .2 Course over ground and ground speed |
| GPZDA 10, 2, 1, 0 or .2 Time and date |

NOTE:  Commands GPRMC and GPVTG are used. GPZDA can be added for computer clock sync.

Other useful commands:

- $JBAUD - Used to configure the baud rate of the receiver.

$JBAUD,R [,OTHER]<CR><LF>

R= 4800, 9600, 19200, 38400, 57600.

- $JSAVE – Command used to save the configuration of the receiver.

$JSAVE<CR><LF>

- $JOFF – This command is used to turn off all data output by the GPS engine.

$JOFF[,OTHER]<CR><LF>

- $JSHOW – Command used to query the receiver for its configuration.

$JSHOW [,SUBSET]<CR><LF>

Using this command without SUBSET will provide a complete response from the receiver.

- $JATT,ROLL To get the roll measurement, the antenna array will need to be installed perpendicular to the vessel’s axis, and send the following command to the Crescent Vector

$JATT,ROLL,YES,<CR><LF>
To return the Crescent Vector to its default mode of producing the pitch measurement, use the following command
$JATT,ROLL,NO,<CR><LF>

- $JASC,GPRMC,10<CR><LF> Recommended minimum specific GNSS (Global Navigation Satellite System) data.

- $JASC,GPVTG,1<CR><LF> Course over ground and ground speed.

- $JASC,GPZDA,1<CR><LF> Time and Date.

- $JASC,GPGLL,1<CR><LF> Geographic position – latitude / longitude.