NOAA/PSD  Ship-based Turbulent Flux Measurement System. 

Second Unit

# Rev 4.8.2

Turbulence Flux System setup on R/V Hi’ialakai.

Turbulence Portable Flux System setup on R/V  Hi’ailakai.

Turbulence Flux System setup on main mast of R/V Hi’ailakai.

The PSD Turbulent Flux System

The PSD Turbulent Flux Measurement System consists of six components:

- A fast turbulence system with ship motion correction mounted on the foremast or on top of a 30’ portable weather station located near the foremast.

The sensors are:

- Ultrasonic anemometer, Gill Wind Master Pro model #1561-PK-020- s/n Y063002.

- Motion-Pak, a six axis motion detector built in house with three Sundstrand rate sensors and three Systron and Donner accelerometer sensors.

- Solar and IR radiation sensors are:

- Radiometers are:

- Two Precision Spectral Pyranometers (PSP)from Appley Inc., both of them installed on a fix platform.

- Two Precision Infrared Radiometers (PIR) from Appley Inc., both of them installed on a fix platform.

Radiometers are mounted on the highest and unobstructed sky-see location of vessel.

- Bulk Meteorology-sensor are:

- Mean T/RH sensor in aspirator, sensor model # HMP235 by Vaisala.

- Near-Skin Surface Temperature (SST).SST measurement is made with a floating (YSI 46040) thermistor deployed off port side with outrigger.

- Optical Scientific Inc.-Optical Precipitation Sensor model ORG-815 DA.

These instruments are mounted on the top of a 30’ high portable weather station near the foremast.

- Li-COR 7500 fast CO2/H2O gas analyzer - mounted on the top of a 30’ high portable weather station located near the foremast.

- A near surface sea surface temperature measurement is made with a floating thermistor deployed off port side with an outrigger.

- Two differential GPS measuring heading, pitch and information.

| Table 1. Instruments and measurements for air-sea interaction, cloud, and precipitation studies. |
| --- |
| Item | System | Measurement |
| 1 | Air-sea flux system | Motion corrected turbulent fluxes |
| 2 | Pyranometer & Pyrgeometer | Downward solar radiative, IR flux |
| 3 | Bulk meteorology | Air Temp, RH, Wind Speed, Rain Rate |
| 4 | Open Path LI-7500 Analyzer | CO2 and H2O |
| 5 | “Sea Snake” | Surface sea-water temperature |
| 6 | GPS | Ship Heading, pith and roll. |

Data set of these instruments are logged, archived and displayed by two computers; one computer dedicated solely to data acquisition while the second computer serves as real time display data and as the platform for daily diagnostics.

All instruments data are sent to the acquisition computer via a selected serial communication port or via a selected wireless radio modem.

Acquisition System Structure

The acquisition System consist of three CR1000 data loggers, five self contained instruments and two Microsoft Windows XP professional computer.

- Campbell CR1000 data logger, slow mean data MEAN1(1Hz)

- Campbell CR1000 data logger, slow mean data MEAN2(1Hz)

- Campbell CR1000 data logger, fast data Ship Motion(10Hz)

- Wind Master Pro Ultrasonic Anemometer Gill (10Hz)

- Open Path CO2/H2O Analyzer- LI-COR 7500 (10Hz)

- Motion Pak- (10Hz)

- Two of Crescent Hemisphere GPS-VS100, outputs GPS, heading, pitch/roll (1Hz and 20Hz).

- PS3DAS03 Microsoft Windows XP professional computer.

- PS3DAS04 Microsoft Windows XP professional computer.

PS3DAS03 computer collect real time data thru 16 serial ports interface PCI device using ports COM3 thru COM10 for flux data acquisition instruments.

COM11 thru COM18 are available for others instruments to be included on the present study.

Data Files

- All logged and archived data file names are formatted as follows:

aaabYYJJJHH_raw.txt

aaa: Instrument identification

b:    Index , default is 0 for one instrument

YY:  Last two digits of year

JJJ:   julian day

HH:  hour

- Time stamp structure is as follows:

Timestamp: mmssmmm

mm: minutes

ss: seconds

mmm: milliseconds

- The data files are logged initially on folder D:\Data. The daily files are moved to a daily folder of the form of D:\Data\<YYJJJ>  at the end of the UTC day.

Archived files have the following format:

- Combined Ship Data - scs0YYJJJhh_raw.txt   (optional depending on ship data logging system)

- Ship Motion data – mot0YYJJJhh_raw.txt

- Met Data logger 1 – me10YYJJJhh_raw.txt

- Met Data logger 2 – me20YYJJJhh_raw.txt

- Temperature and Humidity  - trh0YYJJJhh_raw.txt

- 7500 Licor gas analyzer – lic0YYJJJhh_raw.txt

- GPS Crescent VS100  – gps0YYJJJhh_raw.txt

- Ship Heading and Pitch  Crescent VS100– hed0YYJJJhh_raw.txt

- Ship Position Crescent  VS100– pos0YYJJJhh_raw.txt

- Ship Heading and Roll Crescent  VS100 – rol0YYJJJhh_raw.txt

- Sonic anemometer – son0YYJJJhh_raw.txt

Serial port assignments and data descriptions:

- Serial Port COM1 received GPS data for computer time sync

- Serial Port COM2 reserved.

- Serial Port COM3-Gill Sonic Anemometer.

Data is acquired at 19200 baud rate and a sampling rate of 10Hz.

Data parameters are separate by a comma (,) and file name format is:

son0yyjjjhh_raw.txt.

See below an example of the first and second lines of an hourly Anemometer raw data file.

08/27/2007 18:00 Generic with ASCII Output (SN: 070514)

2545847 B,-001.20,-000.62,+000.22,M,+032.99,00,25

- Timestamp

- Node Address

- U

- V

- W

- M metric/sec units

- Sonic Temperature (degC)

- Status

- Serial Port COM4- LI-7500 Gas Analyzer.

Data is acquired at 19200 baud rate and a sampling rate of 10Hz.

file name format: lic0yyjjjhh_raw.txt

Data parameters are separate by a space.

See below an example of the first two lines of an hourly Licor7500 raw data file.

08/27/2007 18:00 Generic with ASCII Output (SN: 0772)

2540164 249	11.8338	446.337	36.53	83.3

- Timestamp

- Diag

- CO2

- H2O

- Temperature

- Pressure

'Diag' is a diagnostic, 'CO2' and 'H2O' are in units of mmol / m3, Temperature is in C˚, and Pressure is in kPa.

- Serial Port COM5- Mean1data logger- Surface Water Temperature, Rain Rate and Sprayer.

Data is acquired at 9600 baud rate and a sampling rate of 1Hz.

file name format: me10yyjjjhh_raw.txt

Data parameters are separate by a comma (,)

See below an example of the first two lines of an hourly Mean1 raw data file.

09/13/2007 20:00 Mean 1 Datalogger (SN: 7911)

0050811 13.11, 27.03, 0.214, 0.961, 2.258, 4.896, 0.056, 27.52, 18.42, 45.17, 0, 4.998, NAN, 441.2, 0.001,0,0,0.001, 0.01, 0,0,0.046,0.028,0.193,0,0,NAN,0

- Timestamp

- Data logger Battery voltage -average

- Data logger Temperature in Celsius -average

- ORG Rain Gauge Carrier Voltage -average

- ORG Rain Gauge Sensing Voltage -average

- ORG Rain Gauge sensing in mm/h -average

- Aspirator air flow >0 flow is on -average

- Sea snake 4 wire resistance in ohms -average

- Sea snake 4 wire Temperature in C –average

- Data logger Battery voltage –standard deviation

- Data logger Temperature in Celsius –standard deviation

- ORG Rain Gauge Carrier Voltage –standard deviation

- ORG Rain Gauge Sensing Voltage –standard deviation

- ORG Rain Gauge Sensing in mm/h –standard deviation

- Aspirator air flow >0 flow is on –standard deviation

- Sea snake 4 wire resistance in ohms –standard deviation

- Sea snake 4 wire Temperature in C –standard deviation

Note : Only one sea-snake will be used to measure water surface temperature.

- Serial Port COM6-Mean 2 data logger- Solar Radiation PIR and PSP, Surface Water Temperature.

Data is acquired at 9600 baud rate and a sampling rate of 1Hz.

file name format: me20yyjjjhh_raw.txt

Data parameters are separate by a comma (,)

See below an example of the first two lines of an hourly Mean2 raw data file.

09/13/2007 20:00 Mean 2 Datalogger (SN: 7908)

0056349 12.69, 26.41, 0.507, 0.514, 0.452, 2.174,24.33, 23.63,114.2,246.5,2404, 30.11,NAN,491.7, 0.001, 0,0.009000001,0.004,0.021,0.231,0.905,0.397,5.353,26.16,6.461,0.063,NAN,0

1	Timestamp

2	Data logger Battery voltage -Average

3	Data logger Temperature -Average

4	Case1_R-Average - Case1 resistance

5	CaseT_1-Average - Case1 temperature

6	Dome1_R-Average - Dome1 resistance

7	DomeT_1-Average - Dome1 temperature

8	Case2_R-Average - Case2 resistance

9	CaseT_2-Average - Case2 temperature

10	Dome2_R-Average - Dome2 resistance

11	DomeT_2-Average - Dome2 temperature

12	PIR1_V-Average - PIR1 long wave in volts

13	PIR1_wm2-Average - PIR1 long wave in watts/m2

14	PSP1_V-Average - PSP1 short wave in volts

15	PSP1_wm2-Average - PSP1 short wave in w/m2

16	PIR2_V-Average - PIR2 long wave in volts

17	PIR2_wm2-Average - PIR2 long wave in w/m2

18	PSP2_V-Average - PSP2 short wave in volts

19	PSP2_wm2-Average - PSP2 short wave in w/m2

20	Seatemp – Average - sea snake 4 wire Resistance in ohms

21	SeatempC – Average - sea snake 4 wire Temperature in C- (PSD3 formula)

22	SeatempC2 – Average - sea snake 4 wire Temperature in C- (YSI formula)

23	hpa - Average - atmospheric pressure in hPa/mbar

24	Case1_R - Standard Dev - Case1 resistance

25	CaseT_1 - Standard Dev - Case1 temperature

26	Dome1_R - Standard Dev - Dome1 resistance

27	DomeT_1 - Standard Dev - Dome1 temperature

28	Case2_R - Standard Dev - Case2 resistance

29	CaseT_2 - Standard Dev - Case2 temperature

30	Dome2_R - Standard Dev - Dome2 resistance

31	DomeT_2 - Standard Dev - Dome2 temperature

32	PIR1_V - Standard Dev - PIR1 long wave in volts

33	PIR1_wm2 - Standard Dev - PIR1 long wave in watts/m2

34	PSP1_V - Standard Dev - PSP1 short wave in volts

35	PSP1_wm2 - Standard Dev - PSP1 short wave in w/m2

36	PIR2_V - Standard Dev - PIR2 long wave in volts

37	PIR2_wm2 - Standard Dev - PIR2 long wave in w/m2

38	PSP2_V - Standard Dev - PSP2 short wave in volts

39	PSP2_wm2 - Standard Dev - PSP2 short wave in w/m2

40	Seatemp – Standard Dev - sea snake 4 wire Resistance in ohms

41	SeatempC – Standard Dev - sea snake 4 wire Temperature in C- (PSD3 formula)

- SeatempC2 – Standard Dev - sea snake 4 wire Temperature in C- (YSI formula)

43        hpa - Standard Dev - atmospheric pressure in hPa/mbar

- Serial Port COM7- HMP230 T/RH transmitter
Data acquisition at 9600 baud rate with a sampling rate of 1Hz.

file name format: trh1yyjjjhh_raw.txt 
Data parameters are separate by space “ “.
See below an example of the first two lines of an hourly Motion-Pak  raw data file.

5/18/2012 16:00 Vaisala HMP 235 (SN: Y3440138)

4933905  22.604	 27.827

4934732  22.605	 27.840

1 	Time stamp
2 	Temperature

3	Relative Humidity

- Port COM8-Ship motion pitch, roll and yaw motion, measured with

Motion-Pak.

Data is acquired at 19200 baud rate and a sampling rate of 10Hz.

file name format: mot0yyjjjhh_raw.txt

Data parameters are separate by a comma “,”.
See below an example of the first two lines of an hourly Motion-Pak  raw data file.

08/27/2007 18:00 Generic with ASCII Output (SN: 7459)

2535118 -0.03274211,0.9426998,0.001364254,-0.05184167,-0.1364254,-3.847198

- Timestamp

- xrate - Rate-x in volts

- yrate - Rate-y in volts

- zrate - Rate-z in volts

- xaccel - Acceleration-x in volts

- yaccel - Acceleration-y in volts

- zaccel - Acceleration-z in volts

Rates and Accelerations are in Volts.

- Serial Port COM9- Ship Heading and Pitch with Crescent VS100

Data is acquired at 19200 baud rate and a sampling rate of 10Hz.

file name format: hed0yyjjjhh_raw.txt

Data parameters are separate by a comma (,)

See below an example of the first two lines of an hourly Ship heading and pitch raw data file.

07/09/2009 05:00 Generic with ASCII Output (SN: B1)

0000061 $PSAT,HPR,050001.20,250.45,0.89,,N*3D

1	Timestamp

2	$PSAT,HPR  identifier

3	time UTC time (HHMMSS.SS)

4	heading in degrees

5	pitch in degrees

6	roll in degrees

7	type N for GPS, G for gyro

Important note: This port is configured for Heading and Pitch,  therefore line 6 will show nothing.

- Serial Port COM10- Ship position GPS Cresent VS100

Data is acquired at 19200 baud rate and a sampling rate of 10Hz.

There are two kind of GPS file format that will be seen on GPS files, with 3 different GPS codes ($GPRMC), ($GPGGA), ($GPVTG)

file name format: gps0yyjjjhh_raw.txt

Data parameters are separate by a comma (,)

See below an example of the first two lines of an hourly Ship position GPS  raw data file. 06/12/2009 06:00 Generic with ASCII Output (A2)

0000077 $GPRMC,060000.40,A,3251.15194534, N,07956.64350164,W, 0.01,207.50, 120609, 7.3,W,D*3E

0000671 $GPVTG,145.12,T,152.46,M,0.01,N,0.02,K,D*22

$GPRMC coding

- Timestamp

- GPS code $GPRMC

- GPS time (HHMMSS.SS)

- Status ('A' = valid, 'V' = warning)

- Latitude in degrees, minutes and decimal minutes (DDMM.MMMMM)

- N or S latitude

- Longitude in degrees, minutes and decimal minutes(DDDMM.MMMMM)

- W or E longitude

- Speed-over-ground Z.Z in knots

- Track made good, reference to true north

- UTC date DDMMYY

- Magnetic Variation in degrees D.D

- Variation Sense v=E ;V=W

- *CC check sum

$GPVTG coding

- Timestamp

- GPS code $GPVTG

- T

- True course over ground, 000 to 359 degrees

- Magnetic course over ground, 000 to 359 degrees

- M

- Speed over ground, 00.0 to 99.9 knots

- N

- Speed over ground, 00.0 to 99.9 kilometers per hour

- K

- Check sum

- Serial Port COM11- Ship Heading and Roll with Crescent VS100

Data is acquired at 19200 baud rate and a sampling rate of 10Hz.

file name format: rol0yyjjjhh_raw.txt

Data parameters are separate by a comma (,)

See below an example of the first two lines of an hourly Ship heading and roll raw data file.

07/09/2009 08:00 Generic with ASCII Output (SN: B2)

0000061 $PSAT,HPR,080001.20,187.19,,10.53,N*06

1	Timestamp

2	$PSAT,HPR  identifier

3	time UTC time (HHMMSS.SS)

4	heading in degrees

5	pitch in degrees

6	roll in degrees

7	type N for GPS, G for gyro

- Serial Port COM12- All Ship instruments data
Data acquisition at 9600 baud rate with a sampling rate of 1Hz.

file name format: scs0yyjjjhh_raw.txt

1 	Time stamp
2 	To be determent

Important note: This port is configured for heading and roll, therefore line 5 will show nothing.

Data Processing Computers

Data Processing Computers tasks are:

- Computer name: PS03DAS03 Data Archive.

Computer Description: OptiPlex 330 Minitower, Pentium Dual Core E2180/2.0GHz,1M     800FSB

Software:

Data Acquisition Setup software, this is a windows XP service that allows to setup RS232 serial ports, set up archiving location, folder and files archiving names and archive the data with time stamp.

File Manager software, this is a Windows XP service allows to manage files as they are been archived.

Clock synchronization Windows XP service allows sync computer clock with GPS or Network.

- Computer name: PS03DAS04 Data Analysis.

Computer description: OptiPlex 330 Minitower, Pentium Dual Core  E2180/2.0GHz,1M     800FSB

File Manager software, this is a Windows XP service allows to manage files as they are been archived.

MathWorks used as mathematical computation, analysis, visualization, and algorithm developmental programs.

Labview is s a graphical programming environment used to display real time data.

Wireless communication Devices Setup

Six pairs of wireless radio modems are setup to send/receive data between instruments and acquisition computer. They are Sonic1 anemometer, Licor gas analyzer, Motion Pak ship motion, Mean1 data logger, Mean2 Data logger, and Riegl Distance Sensor. All modems are paired and configured on a point to point interface.

Two kind of radio modems are used:

- Data-Linc Group model PLantLinc 5000 with a transmitter frequency of 900 MHz. , refer to User guide for PLR5000.  These modems are used with the Mean1 data logger.

- Microhard System model Spectra 2420 with a transmitter frequency of 2.4 GHz., refer to Operating manual for Spectra 2420. These modems are used with the Sonic, Licor, Mean2, Motion Pak, and Laser.

Radio modem assignments are as follows:

| Parameter | Radio1 | Radio2 | Radio3 | Radio4 | Radio5 | Radio6 |
| --- | --- | --- | --- | --- | --- | --- |
| System | System 2 | System 2 | System 2 | System 2 | System 2 | System 2 |
| Serial # | 010-0014623 | 010-0014628 | 010-00115084 | 010-0015100 | 010-0015087 | 010-0015106 |
| S101 | 0 | 2 | 2 | 0 | 2 | 0 |
| S102 | 38400 | 38400 | 19200 | 19200 | 38400 | 38400 |
| S103 | 230400 | 230400 | 230400 | 230400 | 230400 | 230400 |
| S104 | 10 | 10 | 11 | 11 | 12 | 12 |
| S105 | 1 | 10 | 2 | 1 | 2 | 1 |
| S107 | spyder | spyder | spyder1 | spyder1 | spyder2 | spyder2 |
| S112 | 120 | 120 | 255 | 255 | 255 | 255 |
| S113 | 5 | 5 | 5 | 5 | 5 | 5 |
| S116 | 5 | 5 | 10 | 10 | 10 | 10 |
| S118 | 1 | 1 | 1 | 1 | 1 | 1 |
| S133 | 1 | 1 | 1 | 1 | 1 | 1 |
| S140 | 10 | 1 | 1 | 2 | 1 | 2 |
| Frequency | 2.4GHz | 2.4GHz | 2.4GHz | 2.4GHz | 2.4GHz | 2.4GHz |
| Connected to | COMM 8 of PC | CR1000 s/n 7459 for Motion Pak | COMM 3 of PC | Sonic Interface | COMM 4 of PC | Licor Interface |
| Restrictions | None | None | none | none | none | None |

| Parameter | Radio7 | Radio8 | Radio9 | Radio10 | Radio11 | Radio12 |
| --- | --- | --- | --- | --- | --- | --- |
| System | System 2 | System 2 | System 2 | System 2 | System 2 | System 2 |
| Serial # | 010-0015098 | 010-0015092 | 010-0015085 | 010-0015082 | C090013 | C090012 |
| S101 | 0 | 2 | 0 | 2 | Master=0 | Slave=1 |
| S102 | 9600 | 9600 | 9600 | 9600 | 9600 | 9600 |
| S103 | 230400 | 230400 | 230400 | 230400 | n/a | n/a |
| S104 | 13 | 13 | 14 | 14 | 255 | 255 |
| S105 | 1 | 2 | 1 | 2 | 885-1137 | 885-1200 |
| S107 | spyder3 | spyde3 | spyder4 | spyder4 | n/a | n/a |
| S112 | 255 | 255 | 255 | 255 |  |  |
| S113 | 5 | 5 | 5 | 5 |  |  |
| S116 | 10 | 10 | 10 | 10 |  |  |
| S118 | 1 | 1 | 1 | 1 |  |  |
| S133 | 1 | 1 | 1 | 1 |  |  |
| S140 | 2 | 1 | 2 | 1 | 885-1200 | 885-1137 |
| Frequency | 2.4GHz | 2.4GHz | 2.4GHz | 2.4GHz | 902 MHz | 902 MHz |
| Connected to | COMM 7 of PC | T/RH Interface | COMM 6 of PC | CR1000 s/n 7908 for Mean 2 | COMM 5 of PC | CR1000 s/n 7911 for Mean 1 |
| Restrictions | None | None | none | none | none | None |

NOTE: Radio Modem parameters are configured before installation;  if replacement is needed replace with a pair of radios and program them accordingly to manual.

Note: This figure represents a standard configuration- installation of the portable flux system; it does not necessarily reflect the actual configuration set up on the field.

Note: This figure represents a standard configuration- installation of the portable flux system; it does not necessarily reflect the actual configuration set up on the field.

Sonic Interface Diagram

7500 Licor Interface Diagram

Mean 1 Interface Diagram

Mean 2 Interface Diagram

Motion Pak and Sprayer Interface Diagram

Interface Diagram for T/RH, Licor, Laser

Optical Rain Gauge interface Diagram

Datalogger programs

-Motion Pak Program as 5/15/2012

CR1000 Series Datalogger

'Created by Sergio Pezoa

'Date: May 15, 2012

'Declared Public Variables

PipeLineMode

Public Meas(7)

Alias Meas(1) = xrate

Alias Meas(2) = yrate

Alias Meas(3) = zrate

Alias Meas(4) = xaccel

Alias Meas(5) = yaccel

Alias Meas(6) = zaccel

Alias Meas(7) = temp

Public OutString As String * 200

'Define Data Tables

DataTable (Motion,1,-1)

Sample (7,Meas(),FP2)

EndTable

'Main Program

BeginProg

SerialOpen (ComRS232,38400,0,0,10)

Scan (100,mSec,500,0)

'Motion Pak voltages reading

VoltDiff (Meas(1),1,mV5000,1,True ,0,250,0.001,0.00828)'offset set on 5/14/2012

VoltDiff (Meas(2),1,mV5000,2,True ,0,250,0.001,0.01728)'offset set on 5/14/2012

VoltDiff (Meas(3),1,mV5000,3,True ,0,250,0.001,0.03569)'offset set on 5/14/2012

VoltDiff (Meas(4),1,mV5000,4,True ,0,250,0.001,0.009)'offset set on 5/14/2012

VoltDiff (Meas(5),1,mV5000,5,True ,0,250,0.001,-0.0064)'offset set on 5/14/2012

VoltDiff (Meas(6),1,mV5000,6,True ,0,250,0.001,-0.008)'offset set on 5/14/2012

OutString =xrate+CHR(44)+yrate+CHR(44)+zrate+CHR(44)+xaccel+CHR(44)+yaccel+CHR(44)+

zaccel+CHR(13)+CHR(10)

SerialOut (ComRS232,OutString,"",0,0)

SerialFlush (COMRS232)

CallTable Motion

NextScan

EndProg

-Mean 1 program as 5/15/2012

'Program name: Mean1_second_sys.CR1

'By Sergio Pezoa

'Mean1

'Log sea temp 4 wires, Rain Gauge, Aspirator and Sprayer

'Date written: 5/3/2012

'Modified

'IF "NAN" make it -9999 exception Rain Gauge

Public Flag(1) As Boolean

Public outputflag

Public OutString1 As String * 85

Public OutString2 As String * 85

Public OutValue(17)

Public TimeValue As String * 25

'Declared variables and units

Public Batt_Volt, PTemp

Public Snake_r4W

Public ORG_Sig_v

Public ORG_Car_v

Public SnakeC_4W

Public ORG_mmhr

Public airflow

Public Water_ON

'Declare Constants

Const c1=0.0010295

Const c2=0.0002391

Const c3=0.0000001568

'Constants for ysi serie 46040  c4,c5,c6 -5 to 40 C

Const c4=0.001399937

Const c5=0.000237854

Const c6=0.000000097

'Contant wire resistance for seasnake

'Const r_cond=5.120

Const high=true

Const low=false

'Define Data Table

DataTable(Mean1,outputflag,-1)

DataInterval(0,1,Min,-1)

'Average

Average(1,Batt_Volt,FP2,0)

Average (1,PTemp,FP2,0)

Average(1,ORG_Car_v,FP2,0)

Average(1,ORG_Sig_v,FP2,0)

Average(1,ORG_mmhr,FP2,0)

Average(1, airflow, FP2,0)

Average (1,Snake_r4W,FP2,0)

Average (1,SnakeC_4W,FP2,0)

Average (1,Water_ON,FP2,0)

'standard dev

StdDev(1,Batt_Volt,FP2,0)

StdDev(1,PTemp,FP2,0)

StdDev(1,ORG_Car_v,FP2,0)

StdDev(1,ORG_Sig_v,FP2,0)

StdDev(1,ORG_mmhr,FP2,0)

StdDev(1,airflow,FP2,0)

StdDev(1,Snake_r4W,FP2,0)

StdDev(1,Water_ON,FP2,0)

EndTable

'Main Program

BeginProg

SerialOpen (ComRS232,9600,0,0,10000)

Scan(1,Sec, 3, 0)

If TimeIntoInterval (0,60,Sec) Then Flag(1)=high

'Default Datalogger Battery Voltage measurement Batt_Volt:

Battery(Batt_Volt)

PanelTemp(PTemp,250)

TimeValue=Mean1.Timestamp(1,1)

'Sprayer Controller

If TimeIntoInterval(0,30,Sec) Then Water_ON=0 'set Port 1 to 0

PortSet(9, Water_ON)   'Set port 1 low for time interval

If TimeIntoInterval(1,86370,Sec) Then Water_ON=1 'set PortOn to 0	86370 once every 24 hrs

PortSet(9, Water_ON)   'Set port 1 high for time interval

'-----------------------------------------------------

'sea snake Half Bridge, 4 Wire measurements Seatemp_r4W:

'7H & 7L ,8H & 8L ,E1

'(wire1,2) ==/\/\/\/\/\== (wire3,4)

'CH7- wire1(yellow)to 7H/10k resistor between 7H and 7L/7L to E3

'wire3(green)to GND

'CH8- wire2(red)to 8H/wire4(blue)to 8L

BrHalf4W(Snake_r4W,1,mV2500,mV2500,7,3,1,2500,False,False,0,_60Hz,10011,0)

If StrComp(Snake_r4W,"INF") = 0 Then

SnakeC_4W=-9999

Else

SnakeC_4W = 1/(c4+c5*LN(Snake_r4W)+c6*PWR(LN(Snake_r4W),3))+(-273.16)

EndIf

'Aspirator check if blowing

'2H , E2 used for on/off switch - not used here

'BrHalf(airflow, 1,mV5000, 3,Vx2, 1,100,False, 0,_60Hz,-5,5) -not used here

'Pulse count for Aspirator air flow

'P1

PulseCount (airflow,1,1 ,1,1,1,0)

'Rain Gauge Carrige voltages

'5H &5L

VoltDiff(ORG_Car_v, 1, mV5000,5 ,False, 0, _60Hz, 0.001, 0)

'Rain Gauge sense voltage

'6H & 6L

VoltDiff(ORG_Sig_v, 1,mV2500, 6,False, 0, _60Hz, 0.001,0)

'ORG_mmhr = 0

'If  ORG_Sig_v <= 0.085 Then

'ORG_mmhr= 0

'Else

'rain rate in mm/hr with an standard calibration (500 mm/hr)

ORG_mmhr = ((ORG_Sig_v)^1.87) *25 - 0.15

'EndIf

If TimeIntoInterval(0,60,Sec) Then outputflag=high

If Flag(1)=true Then

GetRecord (OutValue(),Mean1,1)

OutString1=TimeValue+CHR(44)+OutValue(1)+CHR(44)+OutValue(2)+CHR(44)+OutValue(3)+CHR(44)+OutValue(4)+CHR(44)+OutValue(5)+CHR(44)+OutValue(6)+CHR(44)+OutValue(7)+CHR(44)+OutValue(8)+CHR(44)

OutString2=OutValue(9)+CHR(44)+OutValue(10)+CHR(44)+OutValue(11)+CHR(44)+OutValue(12)+CHR(44)+OutValue(13)+CHR(44)+OutValue(14)+CHR(44)+OutValue(15)+CHR(44)+OutValue(16)+CHR(44)+OutValue(17)+CHR(13)+CHR(10)

SerialOut (ComRS232,OutString1+OutString2,"",0,0)

Flag(1)=false

EndIf

CallTable Mean1

NextScan

EndProg

-Mean 2 program as 5/15/2012

'CR1000

'Program name: PSP_PIR_SEC_SYS_VER11.CR1

'Created by Sergio Pezoa

'Mean 2

'Updated 6/18/13

'Logs Sea snake with either 2 or 4 wires, PSP and PIR...

'Declare Variables and Units

'9/28/07 "NAN" changed to -9999

'6/14/13 ....name changed

'Two pyranometer from EPPLEY,

'Two pyrgeometer from EPPLEY,

'Sea Snake, Atmosphoric Pressure

'updated pir's and psp's coeficients 6\14\2013

'changed PSP voltage reading 06/14/2012

'changed mode of reading barometric presure was analog, is serial

'channels used as:

'1H=Eppley PIR1 Case Temp 1

'1L=Eppley PIR1 Dome Temp 3

'2H=Eppley PIR1 thermopile 5

'2L=Eppley PIR1 thermopile 6

'----------

'3H=Eppley PSP1 thermopile

'3L=Eppley PSP1 thermopile

'----------

'4H=Eppley PIR2 Case Temp 1

'4L=Eppley PIR2 Dome Temp 3

'5H=Eppley PIR2 Thermopile 5

'5L=Eppley PIR2 Thermopile 6

'-----------

'6H= Eppley PSP2 Thermopile

'6L= Eppley PSP2 Thermopile

'-----------

'7H=2W sea snake

'7L= to Ground

'----------

'8H= Pressure sensor PTB220 voltage

'8L= to Ground

'-----------

'Declare Variables and Units

Public Batt_Volt,PTemp

Public Case_Res1(2)

Public PIR1_V, PSP1_V

Public Case_Res2(2)

Public PIR2_V, PSP2_V

Public PIR1_wm2, PIR2_wm2

Public PSP1_wm2, PSP2_wm2

Public Case1_R,Dome1_R

Public Case2_R,Dome2_R

Public CaseT_1,CaseT_2

Public DomeT_1,DomeT_2

Public Seatemp_r,Seatemp,SeatempC,SeatempC2

Public pressure_v,pressure

Public hpa As String * 2

'000000

Public outputflag

Public Flag (1) As Boolean

Public OutString As String * 70

Public OutString1 As String * 70

Public OutString2 As String * 70

Public OutString3 As String * 70

Public OutString4 As String * 70

Public OutString5 As String * 70

Public OutString6 As String * 70

Public OutString7 As String * 70

Public count1,TestNum(42)

Public TimeValue

Alias Case_Res1(1) = CaseRes1

Alias Case_Res1(2) = DomeRes1

Alias Case_Res2(1)= CaseRes2

Alias Case_Res2(2)= DomeRes2

'Declare Other Variables

'Declare Constants

'PSP and PIR constants from document 10/23/97 requested by J Hare toEppley

' using YSI 44031

Const c1=0.0010295

Const	c2=0.0002391

Const	c3=0.0000001568

'Constants for YSI serie 46040  c4,c5,c6 do not change

Const c4=0.001399937

Const c5=0.000237854

Const c6=0.000000097

'Contant wire resistance for seasnake

Const r_cond=5.130

'PIR1 s/n 34302F3 calibration on 7/05/2012

Const PIR1_Cnt = 0.00387

'PIR2 s/n 30432F3 calibration on 4/02/2012

Const PIR2_Cnt = 0.00426

'PSP1 s/n 30593F3 calibration on 7/05/2012

Const PSP1_Cnt = 0.00882

'PSP2 s/n 30434F3 calibration on 4/02/2012

Const PSP2_Cnt = 0.00867

Const high=true

Const low = false

'Coefficients fot YSI 44031 PIR1 0 to 40 C

' to calculate thermistor temp

Const c7=0.001017394

Const c8=0.000241046

Const c9=0.000000149

'Define Data Tables

DataTable(Metsec,outputflag,-1)

DataInterval(0,1,Min,-1)

'Average

Average(1,Batt_Volt,FP2,0)

Average (1,PTemp,FP2,0)

'000000

Average (1,Case1_R,IEEE4,0)

Average (1,CaseT_1,FP2,0)

Average (1,Dome1_R,IEEE4,0)

Average (1,DomeT_1,FP2,0)

Average (1,Case2_R,IEEE4,0)

Average (1,CaseT_2,FP2,0)

Average (1,Dome2_R,IEEE4,0)

Average (1,DomeT_2,FP2,0)

'000000

Average (1,PIR1_V,FP2,0)

Average (1,PIR1_wm2,FP2,0)

Average (1,PSP1_V,FP2,0)

Average (1,PSP1_wm2,FP2,0)

'000000

Average (1,PIR2_V,FP2,0)

Average (1,PIR2_wm2,FP2,0)

Average (1,PSP2_V,FP2,0)

Average (1,PSP2_wm2,FP2,0)

'000000

Average (1,Seatemp,FP2,0)

Average (1,SeatempC,FP2,0)

Average (1,SeatempC2,FP2,0)

'000000

'Average (1,pressure_v,FP2,0)

Average (1,hpa,IEEE4,0)

'000000

'Standard Deviation

StdDev (1,Case1_R,IEEE4,False)

StdDev (1,CaseT_1,FP2,False)

StdDev (1,Dome1_R,IEEE4,False)

StdDev (1,DomeT_1,FP2,False)

StdDev (1,Case2_R,IEEE4,False)

StdDev (1,CaseT_2,FP2,False)

StdDev (1,Dome2_R,IEEE4,False)

StdDev (1,DomeT_2,FP2,False)

StdDev (1,PIR1_V,FP2,False)

StdDev (1,PIR1_wm2,FP2,False)

StdDev (1,PSP1_V,FP2,False)

StdDev (1,PSP1_wm2,FP2,False)

StdDev (1,PIR2_V,FP2,False)

StdDev (1,PIR2_wm2,FP2,False)

StdDev (1,PSP2_V,FP2,False)

StdDev (1,PSP2_wm2,FP2,False)

StdDev (1,Seatemp,FP2,False)

StdDev (1,SeatempC,FP2,False)

StdDev (1,SeatempC2,FP2,False)

StdDev (1,hpa,IEEE4,False)

EndTable

'Main Program

BeginProg

SerialOpen (ComRS232,9600,0,0,10000)

SerialOpen (Com1,9600,0,0,50)

Scan(1,Sec,0,0)

TimeValue = Metsec.Timestamp(5,1)

If TimeIntoInterval (0,60,Sec) Then Flag(1) = high

Battery(Batt_Volt)

PanelTemp(PTemp,250)

'Generic Half Bridge Case1, Dome1 measurements from Eppley PIR1:

'1H case resistance with 10K resistor between 1H and E1

'1L dome resistance with 10K resistor between 1L and E1

BrHalf(Case_Res1(),2,Autorange,1,Vx1,2,500,False,0,250,1,0)

Case1_R=10000*(CaseRes1/(1-CaseRes1))

Dome1_R=10000*(DomeRes1/(1-DomeRes1))

CaseT_1 = 1/(c1+c2*LN(Case1_R)+c3*PWR(LN(Case1_R),3))+(-273.16)

DomeT_1 = 1/(c1+c2*LN(Dome1_R)+c3*PWR(LN(Dome1_R),3))+(-273.16)

'Generic Differential Voltage measurements Eppley PIR PIR_V:

'Diff 2H,2L

VoltDiff(PIR1_V,1,mV25,2,False,0,_60Hz,1.0,0)

If StrComp(PIR1_V,"NAN")= 0 Then

PIR1_V = -9991

PIR1_wm2 = -9991

Else

PIR1_wm2=(1/PIR1_Cnt)*PIR1_V

EndIf

'Generic Differential Voltage measurements Eppley PSP1 3H, 3L

'VoltDiff(PSP1_V,1,mV7_5,3,False,0,_60Hz,1.0,0)

VoltDiff(PSP1_V,1,mV25,3,False,0,_60Hz,1.0,0)

If StrComp(PSP1_V,"NAN")= 0 Then

PSP1_V = -9990

PSP1_wm2 = -9991

Else

PSP1_wm2=(1/PSP1_Cnt)*PSP1_V

EndIf

'Generic Half Bridge Case2, Dome2 measurements from Eppley PIR2:

'4H case resistance with 9,987.5 ohms resistor between 1H and E3

'4L dome resistance with 9,992.0 ohms resistor between 1L and E3

BrHalf(Case_Res2(),2,Autorange,7,Vx3,2,500,False,0,250,1,0)

Case2_R=9987.5*(CaseRes2/(1-CaseRes2))

Dome2_R=9992.0*(DomeRes2/(1-DomeRes2))

CaseT_2 = 1/(c1+c2*LN(Case2_R)+c3*PWR(LN(Case2_R),3))+(-273.16)

DomeT_2 = 1/(c1+c2*LN(Dome2_R)+c3*PWR(LN(Dome2_R),3))+(-273.16)

'Generic Differential Voltage measurements Eppley PIR PIR_V:

'Diff 5H,5L

VoltDiff(PIR2_V,1,mV25,5,False,0,_60Hz,1.0,0)

If StrComp(PIR2_V,"NAN")= 0 Then

PIR2_V = -9990

PIR2_wm2 = -9991

Else

PIR2_wm2=(1/PIR2_Cnt)*PIR2_V

EndIf

'Generic Differential Voltage measurements Eppley PSP2 6H, 6L

'VoltDiff(PSP2_V,1,mV7_5,6,False,0,_60Hz,1.0,0)

VoltDiff(PSP2_V,1,mV25,6,False,0,_60Hz,1.0,0)

If StrComp(PSP2_V,"NAN")= 0 Then

PSP2_V = -9990

PSP2_wm2 = -9991

Else

PSP2_wm2=(1/PSP2_Cnt)*PSP2_V

EndIf

'sea snake measuremnet check

'7H & GND with 10K resistor between 7H and E2

BrHalf(Seatemp_r, 1,mV250,13,Vx2, 1,250,False,0,250, 1,0)

Seatemp = 9989*(Seatemp_r/(1.0-Seatemp_r))

If StrComp(Seatemp, 1 ) < 0 Then

SeatempC = -9999

SeatempC2 = -9999

Else

SeatempC = 3915.9/(r_cond+ LN(Seatemp - 10))-273.16

SeatempC2 =1/(c4+c5*LN(Seatemp)+c6*PWR(LN(Seatemp),3))+(-273.16)

EndIf

'Vaisala Presure Measuremnet Com1 no CR no LF from sensor

SerialInBlock (Com1,hpa,20)

'Call Data Tables and Store Data

If TimeIntoInterval(0,60,sec) Then outputflag=high

If Flag (1)=true Then

GetRecord (TestNum(),Metsec,1)

OutString1 =TimeValue+CHR(44)+TestNum(1)+CHR(44)+TestNum(2)+CHR(44)+TestNum(3)+CHR(44)+TestNum(4)+CHR(44)+TestNum(5)+CHR(44)+TestNum(6)+CHR(44)+TestNum(7)+CHR(44)

OutString2 =TestNum(8)+CHR(44)+TestNum(9)+CHR(44)+TestNum(10)+CHR(44)+TestNum(11)+CHR(44)+TestNum(12)+CHR(44)+TestNum(13)+CHR(44)+TestNum(14)+CHR(44)

OutString3 =TestNum(15)+CHR(44)+TestNum(16)+CHR(44)+TestNum(17)+CHR(44)+TestNum(18)+CHR(44)+TestNum(19)+CHR(44)+TestNum(20)+CHR(44)+TestNum(21)+CHR(44)

OutString4 =TestNum(22)+CHR(44)+TestNum(23)+CHR(44)+TestNum(24)+CHR(44)+TestNum(25)+CHR(44)+TestNum(26)+CHR(44)+TestNum(27)+CHR(44)+TestNum(28)+CHR(44)+TestNum(29)+CHR(44)

OutString6 =TestNum(30)+CHR(44)+TestNum(31)+CHR(44)+TestNum(32)+CHR(44)+TestNum(33)+CHR(44)+TestNum(34)+CHR(44)+TestNum(35)+CHR(44)+TestNum(36)+CHR(44)+TestNum(37)+CHR(44)

OutString7 =TestNum(38)+CHR(44)+TestNum(39)+CHR(44)+TestNum(40)+CHR(44)+TestNum(41)+CHR(44)+TestNum(42)+CHR(13)+CHR(10)

SerialOut (ComRS232,OutString1+OutString2+OutString3+OutString4+OutString5+OutString6+OutString7,"",0,0)

SerialFlush (Com1)

Flag (1)=false

EndIf

CallTable(Metsec)

NextScan

EndProg

Table of Sensors Coefficient

| Sensor | Calibration coefficient | Make / Model | Serial Number | Date of calibration |
| --- | --- | --- | --- | --- |
| Precision Spectral Pyranometer | 0.00882 | Eppley / PSP 1 | 30593F3 | July 5, 2012 |
| Precision Spectral Pyranometer | 0.00795 | Eppley / PSP 2 | 32614F3 | Feb 15, 2012 |
| Precision Infrared Radiometer | 0.00426 | Eppley / PIR 1 | 30432F3 | Mar 14, 2012 |
| Precision Infrared Radiometer | 0.00412 | Eppley / PIR 2 | 30558F3 | Feb 15, 2012 |
| Rain Gauge | Offset=0.057mV | OSI/ORG815DA | 08060282 | Apr 30, 2012 |
| Rain Gauge SPARE | Offset=0.093mV | OSI/ORG815DA | 00010149 | Apr 25, 2012 |
| Motion Pak ZERO xrate | 0.00828 | Sundstrand | n/a | May 15,2012 |
| Motion Pak ZERO yrate | 0.01728 | Sundstrand | n/a | May 15,2012 |
| Motion Pak ZERO zrate | 0.03569 | Sundstrand | n/a | May 15,2012 |
| Motion Pak Zero xaccel | 0.009 | Systron & Donner | n/a | May 15,2012 |
| Motion Pak ZERO yaccel | -0.0064 | Systron & Donner | n/a | May 15,2012 |
| Motion Pak ZERO zaccel | -0.008 | Systron & Donner | n/a | May 15,2012 |
| Sea Snake thermistor -5C to 40C | C4=0.001399937 | YSI 46040 series |  | May 9, 2012 |
| Sea Snake thermistor -5C to 40C | C5=0.00237854 | YSI 46040 series |  |  |
| Sea Snake thermistor -5C to 40C | C6=0.000000097 | YSI 46040 series |  |  |
| Temperature / Humidity | n/a | Vaisala/HMT335 | CT10008 | Feb 22, 2012 |
| Class A Digital Barometer | n/a | Vaisala/ PTB220 | A2710003 | checked OK June16, 2013 |
| CO2/H2O Analyzer | Available | Licor / LI-7500 | 75H-0027 | Zeroed for: CO2 Sep 9, 2012 H2O Sep 9, 2012 |

Hemisphere GPS VS100 set up summary

A GPS VS100 system can be setup to measure heading and pitch or heading and roll but not both.

To get heading and pitch measurements, two steps to follow are needed.

- Primary and secondary antenna need to be in line between the bow and the stern of the ship with the primary antenna on the bow side.

- The following command need to be send to the crescent unit

“$JATT,ROLL,NO <CR><LF>”      (default)

To get heading and roll measurements, two setups are needed

- Primary and secondary antenna need to be in line between the port and starboard of the ship with the primary antenna on the port side.

2.   The following command need to be send to the crescent unit

“$JATT,ROLL,YES <CR><LF>”

The “$JASC” command allows the GPS data messages on at a particular update

rate, to be turned on or off. When turning messages on, various update

rates are available, depending on what the requirements are.

This command has the following layout:

$JASC,MSG,R[,OTHER]<CR><LF>

Where “MSG” is the name of the data message and “R” is the message

rate, as shown in the table below. Sending the command without the

optional “,OTHER” data field (without the brackets) will enact a change

on the current port.

Sending a command with a zero value for the “R” field turns off a

message.

| MSG | R (Hz) | Description |
| --- | --- | --- |
| GPGBS 20, 10, 2, 1, 0 or .2 Satellite fault detection used for RAIM |
| GPGGA 20, 10, 2, 1, 0 or .2 GPS fix data |
| GPGLL 20, 10, 2, 1, 0 or .2 Geographic position -latitude/longitude |
| GPGNS 20, 10, 2, 1, 0 or .2 GNSS fix data |
| GPGRS 20, 10, 2, 1, 0 or .2 GNSS range residuals |
| GPGSA 1 or 0 GNSS (Global Navigation Satellite System DOP (and active satellites)) |
| GPGST 1 or 0 GNSS pseudorange error statistics |
| GPGSV 1 or 0 GNSS satellite in view |
| GPRMC 10, 2, 1, 0 or .2 Recommended minimum specific GNSS data |
| GPRRE 1 or 0 Range residual message |
| GPVTG 10, 2, 1, 0 or .2 Course over ground and ground speed |
| GPZDA 10, 2, 1, 0 or .2 Time and date |

NOTE:  Commands GPRMC and GPVTG are used. GPZDA can be added for computer clock sync.

Other useful commands:

- $JBAUD - Used to configure the baud rate of the receiver.

$JBAUD,R [,OTHER]<CR><LF>

R= 4800, 9600, 19200, 38400, 57600.

- $JSAVE – Command used to save the configuration of the receiver.

$JSAVE<CR><LF>

- $JOFF – This command is used to turn off all data output by the GPS engine.

$JOFF[,OTHER]<CR><LF>

- $JSHOW – Command used to query the receiver for its configuration.

$JSHOW [,SUBSET]<CR><LF>

Using this command without SUBSET will provide a complete response from the receiver.

- $JATT,ROLL To get the roll measurement, the antenna array will need to be installed perpendicular to the vessel’s axis, and send the following command to the Crescent Vector

$JATT,ROLL,YES,<CR><LF>
To return the Crescent Vector to its default mode of producing the pitch measurement, use the following command
$JATT,ROLL,NO,<CR><LF>

- $JASC,GPRMC,10<CR><LF> Recommended minimum specific GNSS (Global Navigation Satellite System) data.

- $JASC,GPVTG,1<CR><LF> Course over ground and ground speed.

- $JASC,GPZDA,1<CR><LF> Time and Date.

- $JASC,GPGLL,1<CR><LF> Geographic position – latitude / longitude.

Notes:

- The pressure measurement is in voltage and converted to hPa by

hPa= pressure voltage *120  + 500.

This linear equation is derived from 1100 hPa is equal to 5V and 500 hPa is equal to 0V

There is a correction equation from PTB220 user guide to compensate for height

HPC =  hPa + 0.1176 * h

Hpa = measured pressure

0.1176 = constant

h = height           range is -40…..+40 meters