WHOTS-2022 REPORT – May 25, 2023 - LB

a. Ship Met Instruments

The NOAA R/V Oscar Elton Sette (OES) is outfitted with sensors for air temperature (Ta), relative humidity (RH), barometric pressure (P), wind speed (U), and wind direction (DIR), sea surface temperature (SST) and salinity (SS)

The ship’s Ta and RH is from an RM Young system (model 41382VC) at the top of the main mast (height ~21m ASL).  Pressure is a Vaisala PTB 330 on the bridge (height ~14.6m ASL) with pressure corrected to sea level.  Wind speed and direction are from an RM Young 05103 prop-vane anemometer on the main mast (height ~22.3m ASL).  Height for ship met sensors on the mast is taken as ~ 70 ft. or 21.3 m in this analysis.  Relative wind speed and direction are corrected to true components using ship heading and speed. Sea water temperature and salinity are from a Sea-bird thermosalinograph (SBE45 and SBE38). See photos in figures 1a and 1b.

b. NOAA/PSL Met System for Bulk Fluxes

The PSL Boundary Layer Observations group collected surface meteorology and sea surface temperature data throughout the cruise. The PSL met system deployed on WHOTS 2021 consists of five components (see photos in Figs 2-5):

- Two Vaisala WXT520 weather stations for U, Ta, RH, P and rain rate @ 13.7 m ASL, installed on the bow mast from struts extending  ~1m forward of the mast.

- Two Eppley PSP pyranometers and Two Eppley PIR pyrgeometers located on the starboard aft corner of the pilot house roof.

- Differential GPS unit measuring 1 Hz SOG, COG, heading and pitch angle.

- Floating thermistor sea surface temperature measurement (YSI 46040 thermistor, sea snake) deployed from a davit off the starboard bow.

- Vaisala PTB 220 digital barometer @ 13.1 m ASL on the aft pilot house roof.

These were all logged on a serial data acquisition system in the ship’s main lab.  Measurements were continuous through the duration of the WHOTS.  Raw motion and GPS data are 1 Hz. WXT data are 1 Hz.  Others are averaged to 1 min.

Additional details on PSL instruments are given in Table 1.

c. Notes on PSL and Ship Measurements

Ta / RH / wind / P:  SCS measurements are quite close to observed PSL values for all of these, as shown in Table 2.  PSL deployed duplicate measurements for these parameters.  PTB, WXT1 and WXT2 pressures agree to within 0.1 mb.  RH from the two WXT sensors agrees very well, and we have taken the mean of these as the reference RH in this analysis. Mean Ta from the WXTs is also used as the reference for air temperature, but WXT and the SCS Ta measurements are all very close; less than 0.05°C as seen in Table 2.

The reference wind speed and direction are taken from the WXT sensors, depending on relative wind direction; for relative wind within +/-60° of the bow, we use the mean wind speed and direction from the two WXT sensors; for relative wind directions >60° from the bow, the starboard WXT wind is preferred; and for relative wind directions <–60° the starboard WXT is used.

Radiation Measurements:  The PSP calibrations were last performed in 2019 at the NOAA/GML roof-top radiation calibration facility.  Solar radiometers are calibrated against the sum of two GML reference standards: a sun-tracking pyroheliometer for direct flux and a shaded Eppley model 848 pyranometer for the diffuse component.  A direct comparison was also made to a secondary reference PSP maintained by GML (S/N 73-36).  This calibration can be reverified post-cruise at ESRL.  The two measurements are in close agreement and the mean value from the two pyranometers is used in this report for the reference solar flux.

The PIR radiometers were calibrated at the GML facility in June 2021. The calibration is referenced to an unshaded, unventilated GML PIR maintained by as a secondary standard, referenced to BSRN instruments as the primary standard.  Following the recommendation of GML, the PSL recalibration was determined from a 3-coefficient form of the Albrecht-Cox equation:

LWR= mVc1+σsbTc4-4σsbTd4-Tc4

where mV is the thermopile signal in millivolts, Tc and Td are case and dome temperatures in °K and calibration coefficient c1 is determined from a regression to the reference measurement.

d. Bulk Met Comparisons with WHOTS-17 and WHOTS -18

Comparisons are based on the hourly mean WHOTS real-time data available from the UOP website. Tables 2–8 give the mean, median and standard deviation measurement bias for ship and mooring systems relative to the PSL measurements.  Pressures are adjusted to sea level for all systems.  PSL wind speed is also adjusted for mean flow distortion effects at the ship’s bow: Raw U is increased by 5% and raw V decreased by 15%.  For the comparison with buoy systems, PSL data (U, Ta, RH) are adjusted to z = 3 m with the COARE bulk flux model (v.3.6).  For the ship SCS comparison, PSL U, Ta, and RH are adjusted to 21m.

Measurement comparisons are computed for WHOTS-18 and SCS data collected at station ALOHA over both mooring intercomparison periods: 7/24/22 0800 to 7/25/22 1400 at the WHOTS-17 site and 7/26/22 0800 to 7/27/22 0500 at the WHOTS-18 site.  Statistics are also computed separately over individual intercomparison periods for each mooring.

The moorings have duplicate met sensor and logging systems (systems 1 and 2).  Tables 2-4 present comparison statistics for the ship and WHOTS-18 systems over the two comparison periods.  Tables 5-8 are mooring comparisons for times when the ship was closest each mooring (within 2 nautical miles.).  Figure 6 shows the ship-to-mooring distances over the entire period at station ALOHA.  Figures 7-22 illustrate these intercomparisons in graphical format.

Figures 23-36 are representative time series plots of the raw ship measurements and COARE bulk flux results for one day, 7/27/22.

e. Notes on results

Pressure: Generally close agreement between all systems within 0.5 mb.

Air temperature:  All systems except WHOTS-17 system 2 are about 0.1°C high or low with respect to the mean WXT air temperature.  This is within measurement uncertainty.

Relative humidity:  Agreement between the various systems is ~2.5% (absolute) or better. This is also within measurement uncertainty. Note that WHOTS-18 system 2 is within 0.5% and that the two PSL WXT RH measurements are in very close agreement and within 0.2% (absolute) of the SCS measurement.

Solar radiation: Agreement between PSL and mooring measurements is generally good but scattered, likely due to cloud shading during the intercomparison.

Infrared radiation: Agreement between all mooring systems and the PSL is better than 5 W/m2 overall.

Sea surface temperature:  Agreement between all mooring systems and the PSL sea snake is better than 0.05°C.

Wind: No wind data from WHOTS-17 system 2.  Wind speed agreement with WHOTS-17 system 1 and WHOTS-18 is good and within 0.2 m/s.  True wind direction for all systems is biased by about 7° high or low when compared to PSL. PSL and SCS are within 1° from each other. Wind speed from the SCS anemometer shows a trend of increasing positive bias with increasing speed, which may be a flow distortion effect.

Table 1: PSL instrument details

| Sensor | Calibration coefficient(s) | Make / Model | Serial Number | Date of calibration |
| --- | --- | --- | --- | --- |
| Pyranometer PSP1 | 0.00819 | Eppley PSP 1 | 28210F3 | March 2019 |
| Pyranometer PSP2 | 0.00825 | Eppley PSP 2 | 34290F3 | March 2019 |
| Pyrgeometer PIR1 | 0.00409 | Eppley PIR 1 | 34302F3 | June 2021 |
| Pyrgeometer PIR2 | 0.00423 | Eppley PIR 2 | 30558F3 | June 2021 |
| Sea Snake SST thermistor 0C to 40C | C4=0.00144694 C5=0.00237006 C6=0.000000101 | YSI 46040 series | n/a | n/a |
| PTB Barometer | n/a | Vaisala PTB220 | A2710003 | n/a |
| WXT-1 | n/a | Vaisala WXT-520 | E4020004 | 2021 |
| WXT-2 | n/a | Vaisala WXT-520 | F2620007 | 2021 |

Difference Tables: PSL-measured values minus mooring or ship-measured values over both intercomparison periods.  Only possible for SCS and WHOTS-18, since WHOTS-17 was out of the water during the WHOTS-18 intercomparison.

Table 2: PSL – Ship SCS, entire period at ALOHA

Table 3: PSL – W18-1, entire period at ALOHA

Table 4: PSL – W18-2, entire period at ALOHA

Difference Tables: PSL-measured values minus mooring-measured during the intercomparison periods for each mooring.

Table 5: PSL – W18-1, near buoy

Table 6: PSL – W18-2, near buoy

Table 7: PSL – W17-1, near buoy

Table 8: PSL – W17-2, near buoy

Figure 1a: RV Oscar Elton Sette, main mast.

Figure 1b: RV Oscar Elton Sette, SBE thermosalinograph.

Figure 2: RV Oscar Elton Sette, bow and bridge profile.

Figures 3: Pilot house roof installations: Radiometer mount, mean met and network switch boxes; heading box and power distribution box.

Figure 4: Bow installation for two WXT weather stations.

Figure 5: starboard sea snake boom.

Figure 6: Distance from the ship to the two mooring anchor locations during the comparison period at Station ALOHA.

Figures 7-8: Atmospheric pressure comparison.

Figures 9-10: Air temperature comparison.

Figures 11-12: Relative humidity comparisons.

Figures 13-14: Solar radiation comparisons.

Figures 15-16: Longwave radiation comparisons.

Figures 17-18: Sea surface temperature comparisons.

Figures 19-20: Wind speed comparison.

Figures 21-22: Wind direction comparison.

Figure 23: Temperatures.  Tc and Td are PIR case and dome temperatures.

Figure 24: Relative humidity.

Figure 25: Relative wind speed and direction.

Figure 28: True wind speed and direction.

Figure 29: Atmospheric pressure.

Figure 30: Shortwave (solar) radiation.

Figure 31: Infrared radiation.

Figure 32: Rain rate.

Figure 33: Ship track on Aug 27.  On 7/27 was the return transit to Oahu.

Figure 34: Speed-over-ground and course-over-ground.

Figure 35: Ship heading.

Figure 36: Measured and bulk heat flux components with respect to the ocean.